## Description
General information
Catalogue version: 8 October 2021
Created by Jaroslav Šolc, jsolc@cmi.cz, Czech Metrology Institute
This Excel file contains the catalogue of X-ray energy spectra of electronic brachytherapy (eBT) device Axxent manufactured by Xoft.
The catalogue was compiled from suitable data available in literature and from the data measured or created by the project "18NRM02 PRISM-eBT" (Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices; http://www.ebt-empir.eu/) of the European Metrology Programme for Innovation and Research (EMPIR).
The main goal of the catalogue was to develop a set of harmonised reference X-ray qualities denoted as “eBT equivalent” that would be established for in-laboratory activities by the X-ray tubes usually available for calibration in calibration laboratories, but the tube voltage and the filter are specifically selected to simulate the energy spectrum of the corresponding eBT radiation as closely as possible at the reference point.
In general, the catalogue consists of:
1) X-ray photon fluence eBT spectra measured and/or Monte Carlo simulated in air and/or in water, whenever they were available;
2) eBT-equivalent X-ray photon fluence spectra that match selected eBT spectra. Typically, the eBT-equivalent spectra were calculated by a dedicated software, usually SpekPy software. If available, the eBT-equivalent spectra already realized are also included. Please note that the eBT-equivalent spectra may be site-dependent, based on properties of X-ray generator available (anode angle, inherent filtration) and the distance from anode to measurement point (attenuation in air). Realized eBT-equivalent spectra should be validated by spectrometry.
Specifically, this catalog includes spectra listed below:
Sheet | Description
MC, water, 1 cm, CMI | 1) X-ray spectrum for 50 kVp in 1 cm water depth obtained using Monte Carlo simulation. 2) eBT-equivalent spectrum in 1 cm water depth calculated from 1) with SpekPy software.
MC, water, 1 cm, Hiatt | X-ray spectrum for 50 kVp in 1 cm water depth obtained using Monte Carlo simulation.
Graphical only | X-ray fluence spectrum in air laterally at 50 cm distance measured by NIST presented in graphical form only.

## MC, water, 1 cm, CMI
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
12 July 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Axxent |  |  | This sheet also contains eBT-equivalent spectrum
Manufacturer: | Xoft
Spectra information
origin | Monte Carlo simulation
Monte Carlo (MC) code | MCNP6.2
MC simulation settings | EPRDATA14 photon library
authors | Czech Metrology Institute
location | -
publication | Report D4.2.3 of the EMRP MetrExtRT project, recalculated in 2021 using a slightly modified model based on information available in http://dx.doi.org/10.1118/1.4919280, with new photon data tables and lower bias voltage sigma
DOI | N/A
year of the result | 2013
material | water
source-to-detector distance (cm) | 1
description of source-to-detector distance | laterally from beam axis at the centre of anode cone (same as in the sheet "MC, water, 1 cm, Hiatt")
applicator | none
other information | the device is inserted inside a water cooling sheath
Device setting
nominal bias voltage (kV) | 50
bias voltage sigma (kV) | 0.1
anode current (mA) | N/A
[9 empty] | Visualization of the geometry. Arrow points to the scoring volume | Visualization of points of interactions of electrons with the device tip
eBT-equivalent spectra input parameters
Method | [6 empty] | SpekPy fit
Anode material | [6 empty] | W
Anode angle (degree) | [6 empty] | 21
X-ray generator voltage (kWp) | [6 empty] | 50
Filtration 1 - material and thickness (mm) | Be | [5 empty] | 1
Filtration 2 - material and thickness (mm) | Air | [5 empty] | 964
Filtration 3 - material and thickness (mm) | Al | [5 empty] | 0.825
eBT and eBT-equivalent spectra characteristics
Filtration material, maximum thickness (mm; user-defined) and density (g/cm3) | Al | 1 | 2.7 |  |  |  | 2.7
Emean (in F; keV) |  |  | 28.5888 |  |  |  | 28.1557 | average photon energy
Emean (in Kair; keV) |  |  | 23.3325 |  |  |  | 23.1436 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 21.1705 |  |  |  | 21.19 | energy of monoenergetic photons with the same HVL as the given spectrum
1st HVL in Al - Excel (in Kair; mm) |  |  | 0.87929 |  |  |  | 0.88163 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value; determined by Excel script
1st HVL in Al - Method (in Kair; mm) |  |  | 0.877 |  |  |  | 0.877 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value, determined by SpekPy fit
| eBT spectrum in 1 cm water depth |  |  |  | eBT-equivalent spectrum
applicator diameter (mm) |  |  | none |  |  |  | none
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) |  | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative)
| 7.5 |  |  |  | 1.1
| 7.6 | 7.55 | 0 |  | 1.2 | 1.15 | 0
| 7.7 | 7.65 | 1.9836e-12 |  | 1.3 | 1.25 | 0
| 7.8 | 7.75 | 1.1933e-12 |  | 1.4 | 1.35 | 0
| 7.9 | 7.85 | 2.5612e-12 |  | 1.5 | 1.45 | 0
| 8 | 7.95 | 8.6936e-12 |  | 1.6 | 1.55 | 0
| 8.1 | 8.05 | 7.9579e-12 |  | 1.7 | 1.65 | 0
| 8.2 | 8.15 | 1.2125e-11 |  | 1.8 | 1.75 | 0
| 8.3 | 8.25 | 1.3805e-11 |  | 1.9 | 1.85 | 0
| 8.4 | 8.35 | 7.5838e-11 |  | 2 | 1.95 | 0
| 8.5 | 8.45 | 3.4493e-11 |  | 2.1 | 2.05 | 0
| 8.6 | 8.55 | 6.5635e-11 |  | 2.2 | 2.15 | 0
| 8.7 | 8.65 | 8.0303e-11 |  | 2.3 | 2.25 | 0
| 8.8 | 8.75 | 1.0255e-10 |  | 2.4 | 2.35 | 0
| 8.9 | 8.85 | 1.7143e-10 |  | 2.5 | 2.45 | 0
| 9 | 8.95 | 2.0849e-10 |  | 2.6 | 2.55 | 0
| 9.1 | 9.05 | 2.4716e-10 |  | 2.7 | 2.65 | 0
| 9.2 | 9.15 | 3.7577e-10 |  | 2.8 | 2.75 | 0
| 9.3 | 9.25 | 4.7582e-10 |  | 2.9 | 2.85 | 0
| 9.4 | 9.35 | 6.2542e-10 |  | 3 | 2.95 | 0
| 9.5 | 9.45 | 7.9708e-10 |  | 3.1 | 3.05 | 0
| 9.6 | 9.55 | 1.1453e-09 |  | 3.2 | 3.15 | 0
| 9.7 | 9.65 | 2.8663e-09 |  | 3.3 | 3.25 | 0
| 9.8 | 9.75 | 1.6795e-09 |  | 3.4 | 3.35 | 0
| 9.9 | 9.85 | 2.2518e-09 |  | 3.5 | 3.45 | 0
| 10 | 9.95 | 3.8926e-09 |  | 3.6 | 3.55 | 0
| 10.1 | 10.05 | 2.9771e-09 |  | 3.7 | 3.65 | 0
| 10.2 | 10.15 | 3.8086e-09 |  | 3.8 | 3.75 | 0
| 10.3 | 10.25 | 3.8412e-09 |  | 3.9 | 3.85 | 0
| 10.4 | 10.35 | 4.1615e-09 |  | 4 | 3.95 | 0
| 10.5 | 10.45 | 5.3162e-09 |  | 4.1 | 4.05 | 0
| 10.6 | 10.55 | 6.2444e-09 |  | 4.2 | 4.15 | 0
| 10.7 | 10.65 | 7.7806e-09 |  | 4.3 | 4.25 | 0
| 10.8 | 10.75 | 8.4866e-09 |  | 4.4 | 4.35 | 0
| 10.9 | 10.85 | 1.0216e-08 |  | 4.5 | 4.45 | 0
| 11 | 10.95 | 1.1554e-08 |  | 4.6 | 4.55 | 0
| 11.1 | 11.05 | 1.325e-08 |  | 4.7 | 4.65 | 0
| 11.2 | 11.15 | 1.5481e-08 |  | 4.8 | 4.75 | 0
| 11.3 | 11.25 | 1.7317e-08 |  | 4.9 | 4.85 | 0
| 11.4 | 11.35 | 2.3603e-08 |  | 5 | 4.95 | 0
| 11.5 | 11.45 | 2.0405e-08 |  | 5.1 | 5.05 | 0
| 11.6 | 11.55 | 2.4493e-08 |  | 5.2 | 5.15 | 0
| 11.7 | 11.65 | 2.5773e-08 |  | 5.3 | 5.25 | 0
| 11.8 | 11.75 | 2.8073e-08 |  | 5.4 | 5.35 | 0
| 11.9 | 11.85 | 3.0403e-08 |  | 5.5 | 5.45 | 0
| 12 | 11.95 | 3.3922e-08 |  | 5.6 | 5.55 | 0
| 12.1 | 12.05 | 3.7345e-08 |  | 5.7 | 5.65 | 0
| 12.2 | 12.15 | 3.9054e-08 |  | 5.8 | 5.75 | 0
| 12.3 | 12.25 | 4.4069e-08 |  | 5.9 | 5.85 | 0
| 12.4 | 12.35 | 4.7348e-08 |  | 6 | 5.95 | 0
| 12.5 | 12.45 | 5.183e-08 |  | 6.1 | 6.05 | 0
| 12.6 | 12.55 | 5.721e-08 |  | 6.2 | 6.15 | 0
| 12.7 | 12.65 | 6.1523e-08 |  | 6.3 | 6.25 | 1.597e-17
| 12.8 | 12.75 | 7.0777e-08 |  | 6.4 | 6.35 | 3.194e-17
| 12.9 | 12.85 | 7.1129e-08 |  | 6.5 | 6.45 | 9.582e-17
| 13 | 12.95 | 7.5845e-08 |  | 6.6 | 6.55 | 2.8746e-16
| 13.1 | 13.05 | 8.0158e-08 |  | 6.7 | 6.65 | 7.985e-16
| 13.2 | 13.15 | 8.7454e-08 |  | 6.8 | 6.75 | 2.0601e-15
| 13.3 | 13.25 | 9.5128e-08 |  | 6.9 | 6.85 | 5.0785e-15
| 13.4 | 13.35 | 9.8915e-08 |  | 7 | 6.95 | 1.185e-14
| 13.5 | 13.45 | 1.0714e-07 |  | 7.1 | 7.05 | 2.643e-14
| 13.6 | 13.55 | 1.1288e-07 |  | 7.2 | 7.15 | 5.655e-14
| 13.7 | 13.65 | 1.2117e-07 |  | 7.3 | 7.25 | 1.1637e-13
| 13.8 | 13.75 | 1.296e-07 |  | 7.4 | 7.35 | 2.3042e-13
| 13.9 | 13.85 | 1.3923e-07 |  | 7.5 | 7.45 | 1.7431e-12
| 14 | 13.95 | 1.4509e-07 |  | 7.6 | 7.55 | 8.1222e-13
| 14.1 | 14.05 | 1.6436e-07 |  | 7.7 | 7.65 | 1.4527e-12
| 14.2 | 14.15 | 1.809e-07 |  | 7.8 | 7.75 | 2.5884e-12
| 14.3 | 14.25 | 1.9255e-07 |  | 7.9 | 7.85 | 4.2588e-12
| 14.4 | 14.35 | 2.0657e-07 |  | 8 | 7.95 | 7.0059e-12
| 14.5 | 14.45 | 2.0858e-07 |  | 8.1 | 8.05 | 1.1525e-11
| 14.6 | 14.55 | 2.1896e-07 |  | 8.2 | 8.15 | 1.7656e-11
| 14.7 | 14.65 | 2.3406e-07 |  | 8.3 | 8.25 | 2.7151e-11
| 14.8 | 14.75 | 2.5256e-07 |  | 8.4 | 8.35 | 4.09e-11
| 14.9 | 14.85 | 9.7268e-07 |  | 8.5 | 8.45 | 3.5059e-09
| 15 | 14.95 | 1.658e-06 |  | 8.6 | 8.55 | 8.7667e-11
| 15.1 | 15.05 | 2.2856e-07 |  | 8.7 | 8.65 | 1.2501e-10
| 15.2 | 15.15 | 2.3577e-07 |  | 8.8 | 8.75 | 1.7537e-10
| 15.3 | 15.25 | 2.4427e-07 |  | 8.9 | 8.85 | 4.0041e-10
| 15.4 | 15.35 | 2.5539e-07 |  | 9 | 8.95 | 3.299e-10
| 15.5 | 15.45 | 2.6238e-07 |  | 9.1 | 9.05 | 4.4307e-10
| 15.6 | 15.55 | 2.7421e-07 |  | 9.2 | 9.15 | 5.8763e-10
| 15.7 | 15.65 | 2.8323e-07 |  | 9.3 | 9.25 | 7.7012e-10
| 15.8 | 15.75 | 2.9062e-07 |  | 9.4 | 9.35 | 1.0194e-09
| 15.9 | 15.85 | 2.9967e-07 |  | 9.5 | 9.45 | 1.2778e-09
| 16 | 15.95 | 3.1415e-07 |  | 9.6 | 9.55 | 1.6197e-09
| 16.1 | 16.05 | 3.1479e-07 |  | 9.7 | 9.65 | 5.3654e-09
| 16.2 | 16.15 | 3.2639e-07 |  | 9.8 | 9.75 | 7.0331e-08
| 16.3 | 16.25 | 3.2797e-07 |  | 9.9 | 9.85 | 3.158e-09
| 16.4 | 16.35 | 3.3587e-07 |  | 10 | 9.95 | 1.1901e-08
| 16.5 | 16.45 | 3.4211e-07 |  | 10.1 | 10.05 | 6.1661e-08
| 16.6 | 16.55 | 3.4817e-07 |  | 10.2 | 10.15 | 5.5364e-09
| 16.7 | 16.65 | 5.99e-07 |  | 10.3 | 10.25 | 1.0167e-08
| 16.8 | 16.75 | 6.9403e-07 |  | 10.4 | 10.35 | 6.2323e-09
| 16.9 | 16.85 | 3.5749e-07 |  | 10.5 | 10.45 | 6.305e-09
| 17 | 16.95 | 4.6146e-07 |  | 10.6 | 10.55 | 7.4238e-09
| 17.1 | 17.05 | 2.4863e-07 |  | 10.7 | 10.65 | 8.6874e-09
| 17.2 | 17.15 | 2.5634e-07 |  | 10.8 | 10.75 | 1.011e-08
| 17.3 | 17.25 | 2.6355e-07 |  | 10.9 | 10.85 | 1.1702e-08
| 17.4 | 17.35 | 2.7599e-07 |  | 11 | 10.95 | 1.3471e-08
| 17.5 | 17.45 | 2.8258e-07 |  | 11.1 | 11.05 | 1.8628e-08
| 17.6 | 17.55 | 2.8602e-07 |  | 11.2 | 11.15 | 1.7577e-08
| 17.7 | 17.65 | 2.9099e-07 |  | 11.3 | 11.25 | 2.0077e-08
| 17.8 | 17.75 | 2.9648e-07 |  | 11.4 | 11.35 | 2.2502e-08
| 17.9 | 17.85 | 3.0638e-07 |  | 11.5 | 11.45 | 1.7579e-07
| 18 | 17.95 | 3.0903e-07 |  | 11.6 | 11.55 | 2.9494e-08
| 18.1 | 18.05 | 3.1931e-07 |  | 11.7 | 11.65 | 4.8019e-08
| 18.2 | 18.15 | 3.2326e-07 |  | 11.8 | 11.75 | 5.0738e-08
| 18.3 | 18.25 | 3.3161e-07 |  | 11.9 | 11.85 | 3.3163e-08
| 18.4 | 18.35 | 3.365e-07 |  | 12 | 11.95 | 3.7472e-08
| 18.5 | 18.45 | 3.4175e-07 |  | 12.1 | 12.05 | 4.0334e-08
| 18.6 | 18.55 | 3.4988e-07 |  | 12.2 | 12.15 | 5.2824e-08
| 18.7 | 18.65 | 3.5375e-07 |  | 12.3 | 12.25 | 4.5067e-08
| 18.8 | 18.75 | 3.6396e-07 |  | 12.4 | 12.35 | 4.8828e-08
| 18.9 | 18.85 | 3.6883e-07 |  | 12.5 | 12.45 | 5.3128e-08
| 19 | 18.95 | 3.75e-07 |  | 12.6 | 12.55 | 5.7652e-08
| 19.1 | 19.05 | 3.808e-07 |  | 12.7 | 12.65 | 6.2385e-08
| 19.2 | 19.15 | 3.8521e-07 |  | 12.8 | 12.75 | 6.7325e-08
| 19.3 | 19.25 | 3.967e-07 |  | 12.9 | 12.85 | 7.2476e-08
| 19.4 | 19.35 | 3.9908e-07 |  | 13 | 12.95 | 7.7833e-08
| 19.5 | 19.45 | 3.9964e-07 |  | 13.1 | 13.05 | 8.3394e-08
| 19.6 | 19.55 | 4.1582e-07 |  | 13.2 | 13.15 | 8.9153e-08
| 19.7 | 19.65 | 4.1256e-07 |  | 13.3 | 13.25 | 9.5105e-08
| 19.8 | 19.75 | 4.2241e-07 |  | 13.4 | 13.35 | 1.0124e-07
| 19.9 | 19.85 | 4.2952e-07 |  | 13.5 | 13.45 | 1.0757e-07
| 20 | 19.95 | 4.2649e-07 |  | 13.6 | 13.55 | 1.1406e-07
| 20.1 | 20.05 | 4.3782e-07 |  | 13.7 | 13.65 | 1.2073e-07
| 20.2 | 20.15 | 4.4455e-07 |  | 13.8 | 13.75 | 1.2755e-07
| 20.3 | 20.25 | 4.5504e-07 |  | 13.9 | 13.85 | 1.3453e-07
| 20.4 | 20.35 | 4.5282e-07 |  | 14 | 13.95 | 1.4165e-07
| 20.5 | 20.45 | 4.5952e-07 |  | 14.1 | 14.05 | 1.4891e-07
| 20.6 | 20.55 | 4.6894e-07 |  | 14.2 | 14.15 | 1.5629e-07
| 20.7 | 20.65 | 4.6443e-07 |  | 14.3 | 14.25 | 1.6377e-07
| 20.8 | 20.75 | 4.7321e-07 |  | 14.4 | 14.35 | 1.7136e-07
| 20.9 | 20.85 | 4.7909e-07 |  | 14.5 | 14.45 | 1.7905e-07
| 21 | 20.95 | 4.8449e-07 |  | 14.6 | 14.55 | 1.8683e-07
| 21.1 | 21.05 | 4.8822e-07 |  | 14.7 | 14.65 | 1.947e-07
| 21.2 | 21.15 | 4.9137e-07 |  | 14.8 | 14.75 | 2.0264e-07
| 21.3 | 21.25 | 4.9157e-07 |  | 14.9 | 14.85 | 2.1065e-07
| 21.4 | 21.35 | 5.0169e-07 |  | 15 | 14.95 | 2.1872e-07
| 21.5 | 21.45 | 5.0225e-07 |  | 15.1 | 15.05 | 2.2684e-07
| 21.6 | 21.55 | 5.0342e-07 |  | 15.2 | 15.15 | 2.3497e-07
| 21.7 | 21.65 | 5.1281e-07 |  | 15.3 | 15.25 | 2.431e-07
| 21.8 | 21.75 | 5.172e-07 |  | 15.4 | 15.35 | 2.5126e-07
| 21.9 | 21.85 | 5.1963e-07 |  | 15.5 | 15.45 | 2.5943e-07
| 22 | 21.95 | 6.1205e-07 |  | 15.6 | 15.55 | 2.6761e-07
| 22.1 | 22.05 | 5.2925e-07 |  | 15.7 | 15.65 | 2.7579e-07
| 22.2 | 22.15 | 7.0284e-07 |  | 15.8 | 15.75 | 2.8396e-07
| 22.3 | 22.25 | 5.2392e-07 |  | 15.9 | 15.85 | 2.9212e-07
| 22.4 | 22.35 | 5.321e-07 |  | 16 | 15.95 | 3.0025e-07
| 22.5 | 22.45 | 5.3329e-07 |  | 16.1 | 16.05 | 3.0836e-07
| 22.6 | 22.55 | 5.3548e-07 |  | 16.2 | 16.15 | 3.1644e-07
| 22.7 | 22.65 | 5.4045e-07 |  | 16.3 | 16.25 | 3.2446e-07
| 22.8 | 22.75 | 5.412e-07 |  | 16.4 | 16.35 | 3.3242e-07
| 22.9 | 22.85 | 5.4934e-07 |  | 16.5 | 16.45 | 3.403e-07
| 23 | 22.95 | 5.5247e-07 |  | 16.6 | 16.55 | 3.4813e-07
| 23.1 | 23.05 | 5.518e-07 |  | 16.7 | 16.65 | 3.5588e-07
| 23.2 | 23.15 | 5.5722e-07 |  | 16.8 | 16.75 | 3.6357e-07
| 23.3 | 23.25 | 5.5754e-07 |  | 16.9 | 16.85 | 3.7118e-07
| 23.4 | 23.35 | 5.675e-07 |  | 17 | 16.95 | 3.7871e-07
| 23.5 | 23.45 | 5.6626e-07 |  | 17.1 | 17.05 | 3.8617e-07
| 23.6 | 23.55 | 5.6317e-07 |  | 17.2 | 17.15 | 3.9353e-07
| 23.7 | 23.65 | 5.6331e-07 |  | 17.3 | 17.25 | 4.008e-07
| 23.8 | 23.75 | 5.711e-07 |  | 17.4 | 17.35 | 4.0798e-07
| 23.9 | 23.85 | 5.6822e-07 |  | 17.5 | 17.45 | 4.1506e-07
| 24 | 23.95 | 5.7565e-07 |  | 17.6 | 17.55 | 4.2204e-07
| 24.1 | 24.05 | 5.7738e-07 |  | 17.7 | 17.65 | 4.2885e-07
| 24.2 | 24.15 | 5.8221e-07 |  | 17.8 | 17.75 | 4.3548e-07
| 24.3 | 24.25 | 5.8239e-07 |  | 17.9 | 17.85 | 4.42e-07
| 24.4 | 24.35 | 5.7762e-07 |  | 18 | 17.95 | 4.4841e-07
| 24.5 | 24.45 | 5.7944e-07 |  | 18.1 | 18.05 | 4.5471e-07
| 24.6 | 24.55 | 5.7783e-07 |  | 18.2 | 18.15 | 4.609e-07
| 24.7 | 24.65 | 5.7811e-07 |  | 18.3 | 18.25 | 4.6696e-07
| 24.8 | 24.75 | 5.8261e-07 |  | 18.4 | 18.35 | 4.7291e-07
| 24.9 | 24.85 | 5.9739e-07 |  | 18.5 | 18.45 | 4.7875e-07
| 25 | 24.95 | 6.3615e-07 |  | 18.6 | 18.55 | 4.8446e-07
| 25.1 | 25.05 | 5.8668e-07 |  | 18.7 | 18.65 | 4.9006e-07
| 25.2 | 25.15 | 5.8268e-07 |  | 18.8 | 18.75 | 4.9553e-07
| 25.3 | 25.25 | 5.8872e-07 |  | 18.9 | 18.85 | 5.0088e-07
| 25.4 | 25.35 | 5.8928e-07 |  | 19 | 18.95 | 5.0611e-07
| 25.5 | 25.45 | 6.025e-07 |  | 19.1 | 19.05 | 5.1122e-07
| 25.6 | 25.55 | 5.9282e-07 |  | 19.2 | 19.15 | 5.1621e-07
| 25.7 | 25.65 | 5.9112e-07 |  | 19.3 | 19.25 | 5.2106e-07
| 25.8 | 25.75 | 5.8578e-07 |  | 19.4 | 19.35 | 5.258e-07
| 25.9 | 25.85 | 5.9082e-07 |  | 19.5 | 19.45 | 5.3041e-07
| 26 | 25.95 | 5.9166e-07 |  | 19.6 | 19.55 | 5.3491e-07
| 26.1 | 26.05 | 5.9065e-07 |  | 19.7 | 19.65 | 5.3928e-07
| 26.2 | 26.15 | 5.846e-07 |  | 19.8 | 19.75 | 5.4352e-07
| 26.3 | 26.25 | 5.8669e-07 |  | 19.9 | 19.85 | 5.4764e-07
| 26.4 | 26.35 | 5.8929e-07 |  | 20 | 19.95 | 5.5164e-07
| 26.5 | 26.45 | 5.8839e-07 |  | 20.1 | 20.05 | 5.5553e-07
| 26.6 | 26.55 | 5.9235e-07 |  | 20.2 | 20.15 | 5.5922e-07
| 26.7 | 26.65 | 5.9093e-07 |  | 20.3 | 20.25 | 5.6271e-07
| 26.8 | 26.75 | 5.8827e-07 |  | 20.4 | 20.35 | 5.6609e-07
| 26.9 | 26.85 | 5.9101e-07 |  | 20.5 | 20.45 | 5.6935e-07
| 27 | 26.95 | 5.8986e-07 |  | 20.6 | 20.55 | 5.7251e-07
| 27.1 | 27.05 | 5.8825e-07 |  | 20.7 | 20.65 | 5.7555e-07
| 27.2 | 27.15 | 5.9132e-07 |  | 20.8 | 20.75 | 5.7847e-07
| 27.3 | 27.25 | 5.8582e-07 |  | 20.9 | 20.85 | 5.8129e-07
| 27.4 | 27.35 | 5.8978e-07 |  | 21 | 20.95 | 5.84e-07
| 27.5 | 27.45 | 5.8868e-07 |  | 21.1 | 21.05 | 5.866e-07
| 27.6 | 27.55 | 5.87e-07 |  | 21.2 | 21.15 | 5.891e-07
| 27.7 | 27.65 | 5.9279e-07 |  | 21.3 | 21.25 | 5.9148e-07
| 27.8 | 27.75 | 5.8975e-07 |  | 21.4 | 21.35 | 5.9376e-07
| 27.9 | 27.85 | 5.7847e-07 |  | 21.5 | 21.45 | 5.9594e-07
| 28 | 27.95 | 5.8612e-07 |  | 21.6 | 21.55 | 5.9803e-07
| 28.1 | 28.05 | 5.8725e-07 |  | 21.7 | 21.65 | 6.0001e-07
| 28.2 | 28.15 | 5.8465e-07 |  | 21.8 | 21.75 | 6.0188e-07
| 28.3 | 28.25 | 5.8035e-07 |  | 21.9 | 21.85 | 6.0366e-07
| 28.4 | 28.35 | 5.8196e-07 |  | 22 | 21.95 | 6.0535e-07
| 28.5 | 28.45 | 5.7701e-07 |  | 22.1 | 22.05 | 6.0695e-07
| 28.6 | 28.55 | 5.8092e-07 |  | 22.2 | 22.15 | 6.0837e-07
| 28.7 | 28.65 | 5.7879e-07 |  | 22.3 | 22.25 | 6.0962e-07
| 28.8 | 28.75 | 5.7659e-07 |  | 22.4 | 22.35 | 6.1079e-07
| 28.9 | 28.85 | 5.7326e-07 |  | 22.5 | 22.45 | 6.1187e-07
| 29 | 28.95 | 5.7846e-07 |  | 22.6 | 22.55 | 6.1288e-07
| 29.1 | 29.05 | 5.7032e-07 |  | 22.7 | 22.65 | 6.1379e-07
| 29.2 | 29.15 | 5.7038e-07 |  | 22.8 | 22.75 | 6.1462e-07
| 29.3 | 29.25 | 5.6993e-07 |  | 22.9 | 22.85 | 6.1537e-07
| 29.4 | 29.35 | 5.7033e-07 |  | 23 | 22.95 | 6.1604e-07
| 29.5 | 29.45 | 5.6946e-07 |  | 23.1 | 23.05 | 6.1665e-07
| 29.6 | 29.55 | 5.6764e-07 |  | 23.2 | 23.15 | 6.1717e-07
| 29.7 | 29.65 | 5.658e-07 |  | 23.3 | 23.25 | 6.1761e-07
| 29.8 | 29.75 | 5.6541e-07 |  | 23.4 | 23.35 | 6.1797e-07
| 29.9 | 29.85 | 5.6011e-07 |  | 23.5 | 23.45 | 6.1827e-07
| 30 | 29.95 | 5.5944e-07 |  | 23.6 | 23.55 | 6.185e-07
| 30.1 | 30.05 | 5.5691e-07 |  | 23.7 | 23.65 | 6.1865e-07
| 30.2 | 30.15 | 5.5488e-07 |  | 23.8 | 23.75 | 6.1874e-07
| 30.3 | 30.25 | 5.5887e-07 |  | 23.9 | 23.85 | 6.1875e-07
| 30.4 | 30.35 | 5.5013e-07 |  | 24 | 23.95 | 6.187e-07
| 30.5 | 30.45 | 5.4572e-07 |  | 24.1 | 24.05 | 6.1859e-07
| 30.6 | 30.55 | 5.4833e-07 |  | 24.2 | 24.15 | 6.1842e-07
| 30.7 | 30.65 | 5.4878e-07 |  | 24.3 | 24.25 | 6.1817e-07
| 30.8 | 30.75 | 5.4299e-07 |  | 24.4 | 24.35 | 6.1787e-07
| 30.9 | 30.85 | 5.4406e-07 |  | 24.5 | 24.45 | 6.175e-07
| 31 | 30.95 | 5.4321e-07 |  | 24.6 | 24.55 | 6.1709e-07
| 31.1 | 31.05 | 5.4027e-07 |  | 24.7 | 24.65 | 6.1662e-07
| 31.2 | 31.15 | 5.3773e-07 |  | 24.8 | 24.75 | 6.1608e-07
| 31.3 | 31.25 | 5.3507e-07 |  | 24.9 | 24.85 | 6.1548e-07
| 31.4 | 31.35 | 5.3401e-07 |  | 25 | 24.95 | 6.1484e-07
| 31.5 | 31.45 | 5.3028e-07 |  | 25.1 | 25.05 | 6.1415e-07
| 31.6 | 31.55 | 5.3207e-07 |  | 25.2 | 25.15 | 6.1335e-07
| 31.7 | 31.65 | 5.2409e-07 |  | 25.3 | 25.25 | 6.1245e-07
| 31.8 | 31.75 | 5.2331e-07 |  | 25.4 | 25.35 | 6.115e-07
| 31.9 | 31.85 | 5.2038e-07 |  | 25.5 | 25.45 | 6.105e-07
| 32 | 31.95 | 5.1904e-07 |  | 25.6 | 25.55 | 6.0947e-07
| 32.1 | 32.05 | 5.1664e-07 |  | 25.7 | 25.65 | 6.0839e-07
| 32.2 | 32.15 | 5.1415e-07 |  | 25.8 | 25.75 | 6.0726e-07
| 32.3 | 32.25 | 5.095e-07 |  | 25.9 | 25.85 | 6.0609e-07
| 32.4 | 32.35 | 5.1221e-07 |  | 26 | 25.95 | 6.0487e-07
| 32.5 | 32.45 | 5.1233e-07 |  | 26.1 | 26.05 | 6.0363e-07
| 32.6 | 32.55 | 5.0785e-07 |  | 26.2 | 26.15 | 6.0235e-07
| 32.7 | 32.65 | 5.0181e-07 |  | 26.3 | 26.25 | 6.0101e-07
| 32.8 | 32.75 | 4.9524e-07 |  | 26.4 | 26.35 | 5.9964e-07
| 32.9 | 32.85 | 4.965e-07 |  | 26.5 | 26.45 | 5.9823e-07
| 33 | 32.95 | 4.9355e-07 |  | 26.6 | 26.55 | 5.968e-07
| 33.1 | 33.05 | 4.8692e-07 |  | 26.7 | 26.65 | 5.9532e-07
| 33.2 | 33.15 | 4.9166e-07 |  | 26.8 | 26.75 | 5.938e-07
| 33.3 | 33.25 | 4.8559e-07 |  | 26.9 | 26.85 | 5.9224e-07
| 33.4 | 33.35 | 4.89e-07 |  | 27 | 26.95 | 5.9066e-07
| 33.5 | 33.45 | 4.816e-07 |  | 27.1 | 27.05 | 5.8906e-07
| 33.6 | 33.55 | 4.8007e-07 |  | 27.2 | 27.15 | 5.8741e-07
| 33.7 | 33.65 | 4.7645e-07 |  | 27.3 | 27.25 | 5.8573e-07
| 33.8 | 33.75 | 4.7292e-07 |  | 27.4 | 27.35 | 5.8402e-07
| 33.9 | 33.85 | 4.7541e-07 |  | 27.5 | 27.45 | 5.8228e-07
| 34 | 33.95 | 4.714e-07 |  | 27.6 | 27.55 | 5.8046e-07
| 34.1 | 34.05 | 4.6783e-07 |  | 27.7 | 27.65 | 5.7857e-07
| 34.2 | 34.15 | 4.6585e-07 |  | 27.8 | 27.75 | 5.7664e-07
| 34.3 | 34.25 | 4.6204e-07 |  | 27.9 | 27.85 | 5.7469e-07
| 34.4 | 34.35 | 4.5531e-07 |  | 28 | 27.95 | 5.7272e-07
| 34.5 | 34.45 | 4.5777e-07 |  | 28.1 | 28.05 | 5.7075e-07
| 34.6 | 34.55 | 4.524e-07 |  | 28.2 | 28.15 | 5.6874e-07
| 34.7 | 34.65 | 4.4856e-07 |  | 28.3 | 28.25 | 5.667e-07
| 34.8 | 34.75 | 4.4802e-07 |  | 28.4 | 28.35 | 5.6463e-07
| 34.9 | 34.85 | 4.4659e-07 |  | 28.5 | 28.45 | 5.6255e-07
| 35 | 34.95 | 4.4215e-07 |  | 28.6 | 28.55 | 5.6047e-07
| 35.1 | 35.05 | 4.3867e-07 |  | 28.7 | 28.65 | 5.5835e-07
| 35.2 | 35.15 | 4.3529e-07 |  | 28.8 | 28.75 | 5.5621e-07
| 35.3 | 35.25 | 4.3201e-07 |  | 28.9 | 28.85 | 5.5404e-07
| 35.4 | 35.35 | 4.3e-07 |  | 29 | 28.95 | 5.5186e-07
| 35.5 | 35.45 | 4.2644e-07 |  | 29.1 | 29.05 | 5.4969e-07
| 35.6 | 35.55 | 4.2329e-07 |  | 29.2 | 29.15 | 5.4748e-07
| 35.7 | 35.65 | 4.2024e-07 |  | 29.3 | 29.25 | 5.4524e-07
| 35.8 | 35.75 | 4.2186e-07 |  | 29.4 | 29.35 | 5.4299e-07
| 35.9 | 35.85 | 4.174e-07 |  | 29.5 | 29.45 | 5.4072e-07
| 36 | 35.95 | 4.1197e-07 |  | 29.6 | 29.55 | 5.3846e-07
| 36.1 | 36.05 | 4.0916e-07 |  | 29.7 | 29.65 | 5.3617e-07
| 36.2 | 36.15 | 4.0379e-07 |  | 29.8 | 29.75 | 5.3386e-07
| 36.3 | 36.25 | 3.9913e-07 |  | 29.9 | 29.85 | 5.3153e-07
| 36.4 | 36.35 | 4.0231e-07 |  | 30 | 29.95 | 5.292e-07
| 36.5 | 36.45 | 4.0122e-07 |  | 30.1 | 30.05 | 5.2686e-07
| 36.6 | 36.55 | 3.9224e-07 |  | 30.2 | 30.15 | 5.2447e-07
| 36.7 | 36.65 | 3.9169e-07 |  | 30.3 | 30.25 | 5.2201e-07
| 36.8 | 36.75 | 3.9276e-07 |  | 30.4 | 30.35 | 5.1955e-07
| 36.9 | 36.85 | 3.8373e-07 |  | 30.5 | 30.45 | 5.1708e-07
| 37 | 36.95 | 3.8546e-07 |  | 30.6 | 30.55 | 5.1462e-07
| 37.1 | 37.05 | 3.7989e-07 |  | 30.7 | 30.65 | 5.1213e-07
| 37.2 | 37.15 | 3.7828e-07 |  | 30.8 | 30.75 | 5.0963e-07
| 37.3 | 37.25 | 3.7028e-07 |  | 30.9 | 30.85 | 5.0712e-07
| 37.4 | 37.35 | 3.697e-07 |  | 31 | 30.95 | 5.0461e-07
| 37.5 | 37.45 | 3.6801e-07 |  | 31.1 | 31.05 | 5.0211e-07
| 37.6 | 37.55 | 3.6528e-07 |  | 31.2 | 31.15 | 4.9958e-07
| 37.7 | 37.65 | 3.6667e-07 |  | 31.3 | 31.25 | 4.9704e-07
| 37.8 | 37.75 | 3.5818e-07 |  | 31.4 | 31.35 | 4.9449e-07
| 37.9 | 37.85 | 3.5598e-07 |  | 31.5 | 31.45 | 4.9194e-07
| 38 | 37.95 | 3.5016e-07 |  | 31.6 | 31.55 | 4.8941e-07
| 38.1 | 38.05 | 3.4847e-07 |  | 31.7 | 31.65 | 4.8685e-07
| 38.2 | 38.15 | 3.4958e-07 |  | 31.8 | 31.75 | 4.8427e-07
| 38.3 | 38.25 | 3.4704e-07 |  | 31.9 | 31.85 | 4.817e-07
| 38.4 | 38.35 | 3.438e-07 |  | 32 | 31.95 | 4.7912e-07
| 38.5 | 38.45 | 3.3891e-07 |  | 32.1 | 32.05 | 4.7656e-07
| 38.6 | 38.55 | 3.3706e-07 |  | 32.2 | 32.15 | 4.7397e-07
| 38.7 | 38.65 | 3.3172e-07 |  | 32.3 | 32.25 | 4.7138e-07
| 38.8 | 38.75 | 3.3201e-07 |  | 32.4 | 32.35 | 4.6878e-07
| 38.9 | 38.85 | 3.1986e-07 |  | 32.5 | 32.45 | 4.6618e-07
| 39 | 38.95 | 3.2388e-07 |  | 32.6 | 32.55 | 4.6359e-07
| 39.1 | 39.05 | 3.2147e-07 |  | 32.7 | 32.65 | 4.6099e-07
| 39.2 | 39.15 | 3.162e-07 |  | 32.8 | 32.75 | 4.5837e-07
| 39.3 | 39.25 | 3.1149e-07 |  | 32.9 | 32.85 | 4.5575e-07
| 39.4 | 39.35 | 3.0851e-07 |  | 33 | 32.95 | 4.5314e-07
| 39.5 | 39.45 | 3.0753e-07 |  | 33.1 | 33.05 | 4.5053e-07
| 39.6 | 39.55 | 3.0342e-07 |  | 33.2 | 33.15 | 4.4791e-07
| 39.7 | 39.65 | 2.9903e-07 |  | 33.3 | 33.25 | 4.4527e-07
| 39.8 | 39.75 | 2.9473e-07 |  | 33.4 | 33.35 | 4.4265e-07
| 39.9 | 39.85 | 2.9272e-07 |  | 33.5 | 33.45 | 4.4002e-07
| 40 | 39.95 | 2.8827e-07 |  | 33.6 | 33.55 | 4.374e-07
| 40.1 | 40.05 | 2.9069e-07 |  | 33.7 | 33.65 | 4.3476e-07
| 40.2 | 40.15 | 2.8463e-07 |  | 33.8 | 33.75 | 4.3212e-07
| 40.3 | 40.25 | 2.8281e-07 |  | 33.9 | 33.85 | 4.2949e-07
| 40.4 | 40.35 | 2.7681e-07 |  | 34 | 33.95 | 4.2685e-07
| 40.5 | 40.45 | 2.7364e-07 |  | 34.1 | 34.05 | 4.2422e-07
| 40.6 | 40.55 | 2.6912e-07 |  | 34.2 | 34.15 | 4.2158e-07
| 40.7 | 40.65 | 2.6947e-07 |  | 34.3 | 34.25 | 4.1893e-07
| 40.8 | 40.75 | 2.6572e-07 |  | 34.4 | 34.35 | 4.1628e-07
| 40.9 | 40.85 | 2.5964e-07 |  | 34.5 | 34.45 | 4.1364e-07
| 41 | 40.95 | 2.593e-07 |  | 34.6 | 34.55 | 4.1101e-07
| 41.1 | 41.05 | 2.5606e-07 |  | 34.7 | 34.65 | 4.0836e-07
| 41.2 | 41.15 | 2.5493e-07 |  | 34.8 | 34.75 | 4.0571e-07
| 41.3 | 41.25 | 2.5076e-07 |  | 34.9 | 34.85 | 4.0306e-07
| 41.4 | 41.35 | 2.4703e-07 |  | 35 | 34.95 | 4.0042e-07
| 41.5 | 41.45 | 2.4397e-07 |  | 35.1 | 35.05 | 3.9779e-07
| 41.6 | 41.55 | 2.3953e-07 |  | 35.2 | 35.15 | 3.9511e-07
| 41.7 | 41.65 | 2.348e-07 |  | 35.3 | 35.25 | 3.9241e-07
| 41.8 | 41.75 | 2.3539e-07 |  | 35.4 | 35.35 | 3.8972e-07
| 41.9 | 41.85 | 2.3252e-07 |  | 35.5 | 35.45 | 3.8703e-07
| 42 | 41.95 | 2.2723e-07 |  | 35.6 | 35.55 | 3.8435e-07
| 42.1 | 42.05 | 2.2383e-07 |  | 35.7 | 35.65 | 3.8165e-07
| 42.2 | 42.15 | 2.1946e-07 |  | 35.8 | 35.75 | 3.7896e-07
| 42.3 | 42.25 | 2.1937e-07 |  | 35.9 | 35.85 | 3.7627e-07
| 42.4 | 42.35 | 2.1682e-07 |  | 36 | 35.95 | 3.7359e-07
| 42.5 | 42.45 | 2.1192e-07 |  | 36.1 | 36.05 | 3.7092e-07
| 42.6 | 42.55 | 2.0896e-07 |  | 36.2 | 36.15 | 3.6824e-07
| 42.7 | 42.65 | 2.045e-07 |  | 36.3 | 36.25 | 3.6556e-07
| 42.8 | 42.75 | 2.0232e-07 |  | 36.4 | 36.35 | 3.6288e-07
| 42.9 | 42.85 | 1.9928e-07 |  | 36.5 | 36.45 | 3.6022e-07
| 43 | 42.95 | 1.9738e-07 |  | 36.6 | 36.55 | 3.5756e-07
| 43.1 | 43.05 | 1.9344e-07 |  | 36.7 | 36.65 | 3.5488e-07
| 43.2 | 43.15 | 1.9288e-07 |  | 36.8 | 36.75 | 3.5222e-07
| 43.3 | 43.25 | 1.9031e-07 |  | 36.9 | 36.85 | 3.4955e-07
| 43.4 | 43.35 | 1.868e-07 |  | 37 | 36.95 | 3.469e-07
| 43.5 | 43.45 | 1.826e-07 |  | 37.1 | 37.05 | 3.4425e-07
| 43.6 | 43.55 | 1.8054e-07 |  | 37.2 | 37.15 | 3.4159e-07
| 43.7 | 43.65 | 1.7747e-07 |  | 37.3 | 37.25 | 3.3894e-07
| 43.8 | 43.75 | 1.737e-07 |  | 37.4 | 37.35 | 3.3629e-07
| 43.9 | 43.85 | 1.7136e-07 |  | 37.5 | 37.45 | 3.3364e-07
| 44 | 43.95 | 1.6888e-07 |  | 37.6 | 37.55 | 3.3101e-07
| 44.1 | 44.05 | 1.6592e-07 |  | 37.7 | 37.65 | 3.2836e-07
| 44.2 | 44.15 | 1.6187e-07 |  | 37.8 | 37.75 | 3.2572e-07
| 44.3 | 44.25 | 1.6235e-07 |  | 37.9 | 37.85 | 3.2308e-07
| 44.4 | 44.35 | 1.5851e-07 |  | 38 | 37.95 | 3.2045e-07
| 44.5 | 44.45 | 1.5539e-07 |  | 38.1 | 38.05 | 3.1783e-07
| 44.6 | 44.55 | 1.5466e-07 |  | 38.2 | 38.15 | 3.1519e-07
| 44.7 | 44.65 | 1.4918e-07 |  | 38.3 | 38.25 | 3.1256e-07
| 44.8 | 44.75 | 1.4811e-07 |  | 38.4 | 38.35 | 3.0994e-07
| 44.9 | 44.85 | 1.4696e-07 |  | 38.5 | 38.45 | 3.0732e-07
| 45 | 44.95 | 1.4481e-07 |  | 38.6 | 38.55 | 3.0471e-07
| 45.1 | 45.05 | 1.4094e-07 |  | 38.7 | 38.65 | 3.0209e-07
| 45.2 | 45.15 | 1.3907e-07 |  | 38.8 | 38.75 | 2.9947e-07
| 45.3 | 45.25 | 1.3587e-07 |  | 38.9 | 38.85 | 2.9686e-07
| 45.4 | 45.35 | 1.3182e-07 |  | 39 | 38.95 | 2.9426e-07
| 45.5 | 45.45 | 1.3079e-07 |  | 39.1 | 39.05 | 2.9166e-07
| 45.6 | 45.55 | 1.2866e-07 |  | 39.2 | 39.15 | 2.8905e-07
| 45.7 | 45.65 | 1.2401e-07 |  | 39.3 | 39.25 | 2.8644e-07
| 45.8 | 45.75 | 1.2444e-07 |  | 39.4 | 39.35 | 2.8384e-07
| 45.9 | 45.85 | 1.2059e-07 |  | 39.5 | 39.45 | 2.8124e-07
| 46 | 45.95 | 1.1904e-07 |  | 39.6 | 39.55 | 2.7865e-07
| 46.1 | 46.05 | 1.159e-07 |  | 39.7 | 39.65 | 2.7606e-07
| 46.2 | 46.15 | 1.1471e-07 |  | 39.8 | 39.75 | 2.7346e-07
| 46.3 | 46.25 | 1.1063e-07 |  | 39.9 | 39.85 | 2.7088e-07
| 46.4 | 46.35 | 1.0826e-07 |  | 40 | 39.95 | 2.6829e-07
| 46.5 | 46.45 | 1.0597e-07 |  | 40.1 | 40.05 | 2.6572e-07
| 46.6 | 46.55 | 1.0227e-07 |  | 40.2 | 40.15 | 2.6312e-07
| 46.7 | 46.65 | 1.0128e-07 |  | 40.3 | 40.25 | 2.6052e-07
| 46.8 | 46.75 | 9.7696e-08 |  | 40.4 | 40.35 | 2.5792e-07
| 46.9 | 46.85 | 9.5785e-08 |  | 40.5 | 40.45 | 2.5533e-07
| 47 | 46.95 | 9.2893e-08 |  | 40.6 | 40.55 | 2.5275e-07
| 47.1 | 47.05 | 9.0602e-08 |  | 40.7 | 40.65 | 2.5016e-07
| 47.2 | 47.15 | 8.7507e-08 |  | 40.8 | 40.75 | 2.4757e-07
| 47.3 | 47.25 | 8.4028e-08 |  | 40.9 | 40.85 | 2.4499e-07
| 47.4 | 47.35 | 8.2549e-08 |  | 41 | 40.95 | 2.4241e-07
| 47.5 | 47.45 | 7.899e-08 |  | 41.1 | 41.05 | 2.3984e-07
| 47.6 | 47.55 | 7.7503e-08 |  | 41.2 | 41.15 | 2.3726e-07
| 47.7 | 47.65 | 7.3868e-08 |  | 41.3 | 41.25 | 2.3468e-07
| 47.8 | 47.75 | 7.1399e-08 |  | 41.4 | 41.35 | 2.321e-07
| 47.9 | 47.85 | 6.7019e-08 |  | 41.5 | 41.45 | 2.2954e-07
| 48 | 47.95 | 6.7549e-08 |  | 41.6 | 41.55 | 2.2698e-07
| 48.1 | 48.05 | 6.3889e-08 |  | 41.7 | 41.65 | 2.2441e-07
| 48.2 | 48.15 | 6.1147e-08 |  | 41.8 | 41.75 | 2.2184e-07
| 48.3 | 48.25 | 5.7404e-08 |  | 41.9 | 41.85 | 2.1928e-07
| 48.4 | 48.35 | 5.5116e-08 |  | 42 | 41.95 | 2.1673e-07
| 48.5 | 48.45 | 5.2201e-08 |  | 42.1 | 42.05 | 2.1418e-07
| 48.6 | 48.55 | 5.0343e-08 |  | 42.2 | 42.15 | 2.1163e-07
| 48.7 | 48.65 | 4.696e-08 |  | 42.3 | 42.25 | 2.0907e-07
| 48.8 | 48.75 | 4.4641e-08 |  | 42.4 | 42.35 | 2.0652e-07
| 48.9 | 48.85 | 4.0383e-08 |  | 42.5 | 42.45 | 2.0398e-07
| 49 | 48.95 | 3.8389e-08 |  | 42.6 | 42.55 | 2.0144e-07
| 49.1 | 49.05 | 3.5246e-08 |  | 42.7 | 42.65 | 1.9889e-07
| 49.2 | 49.15 | 3.2269e-08 |  | 42.8 | 42.75 | 1.9635e-07
| 49.3 | 49.25 | 2.9214e-08 |  | 42.9 | 42.85 | 1.938e-07
| 49.4 | 49.35 | 2.5427e-08 |  | 43 | 42.95 | 1.9127e-07
| 49.5 | 49.45 | 2.271e-08 |  | 43.1 | 43.05 | 1.8874e-07
| 49.6 | 49.55 | 1.975e-08 |  | 43.2 | 43.15 | 1.862e-07
| 49.7 | 49.65 | 1.6414e-08 |  | 43.3 | 43.25 | 1.8366e-07
| 49.8 | 49.75 | 1.2762e-08 |  | 43.4 | 43.35 | 1.8112e-07
| 49.9 | 49.85 | 9.2965e-09 |  | 43.5 | 43.45 | 1.7859e-07
| 50 | 49.95 | 5.318e-09 |  | 43.6 | 43.55 | 1.7607e-07
| 50.1 | 50.05 | 1.3366e-09 |  | 43.7 | 43.65 | 1.7354e-07
| 50.2 | 50.15 | 1.3324e-10 |  | 43.8 | 43.75 | 1.71e-07
| 50.3 | 50.25 | 0 |  | 43.9 | 43.85 | 1.6846e-07
[5 empty] | 44 | 43.95 | 1.6594e-07
[5 empty] | 44.1 | 44.05 | 1.6342e-07
[5 empty] | 44.2 | 44.15 | 1.609e-07
[5 empty] | 44.3 | 44.25 | 1.5836e-07
[5 empty] | 44.4 | 44.35 | 1.5583e-07
[5 empty] | 44.5 | 44.45 | 1.5331e-07
[5 empty] | 44.6 | 44.55 | 1.5079e-07
[5 empty] | 44.7 | 44.65 | 1.4826e-07
[5 empty] | 44.8 | 44.75 | 1.4573e-07
[5 empty] | 44.9 | 44.85 | 1.4318e-07
[5 empty] | 45 | 44.95 | 1.4066e-07
[5 empty] | 45.1 | 45.05 | 1.3814e-07
[5 empty] | 45.2 | 45.15 | 1.356e-07
[5 empty] | 45.3 | 45.25 | 1.3305e-07
[5 empty] | 45.4 | 45.35 | 1.3049e-07
[5 empty] | 45.5 | 45.45 | 1.2795e-07
[5 empty] | 45.6 | 45.55 | 1.2541e-07
[5 empty] | 45.7 | 45.65 | 1.2286e-07
[5 empty] | 45.8 | 45.75 | 1.2029e-07
[5 empty] | 45.9 | 45.85 | 1.1771e-07
[5 empty] | 46 | 45.95 | 1.1516e-07
[5 empty] | 46.1 | 46.05 | 1.126e-07
[5 empty] | 46.2 | 46.15 | 1.1003e-07
[5 empty] | 46.3 | 46.25 | 1.0744e-07
[5 empty] | 46.4 | 46.35 | 1.0484e-07
[5 empty] | 46.5 | 46.45 | 1.0226e-07
[5 empty] | 46.6 | 46.55 | 9.9683e-08
[5 empty] | 46.7 | 46.65 | 9.7085e-08
[5 empty] | 46.8 | 46.75 | 9.4462e-08
[5 empty] | 46.9 | 46.85 | 9.1835e-08
[5 empty] | 47 | 46.95 | 8.9229e-08
[5 empty] | 47.1 | 47.05 | 8.6616e-08
[5 empty] | 47.2 | 47.15 | 8.3985e-08
[5 empty] | 47.3 | 47.25 | 8.1327e-08
[5 empty] | 47.4 | 47.35 | 7.8665e-08
[5 empty] | 47.5 | 47.45 | 7.6024e-08
[5 empty] | 47.6 | 47.55 | 7.3376e-08
[5 empty] | 47.7 | 47.65 | 7.0707e-08
[5 empty] | 47.8 | 47.75 | 6.8013e-08
[5 empty] | 47.9 | 47.85 | 6.5312e-08
[5 empty] | 48 | 47.95 | 6.2634e-08
[5 empty] | 48.1 | 48.05 | 5.9948e-08
[5 empty] | 48.2 | 48.15 | 5.7242e-08
[5 empty] | 48.3 | 48.25 | 5.4512e-08
[5 empty] | 48.4 | 48.35 | 5.1773e-08
[5 empty] | 48.5 | 48.45 | 4.9058e-08
[5 empty] | 48.6 | 48.55 | 4.6335e-08
[5 empty] | 48.7 | 48.65 | 4.3604e-08
[5 empty] | 48.8 | 48.75 | 4.085e-08
[5 empty] | 48.9 | 48.85 | 3.8079e-08
[5 empty] | 49 | 48.95 | 3.5332e-08
[5 empty] | 49.1 | 49.05 | 3.2577e-08
[5 empty] | 49.2 | 49.15 | 2.9852e-08
[5 empty] | 49.3 | 49.25 | 2.7099e-08
[5 empty] | 49.4 | 49.35 | 2.4291e-08
[5 empty] | 49.5 | 49.45 | 2.1487e-08
[5 empty] | 49.6 | 49.55 | 1.8671e-08
[5 empty] | 49.7 | 49.65 | 1.5963e-08
[5 empty] | 49.8 | 49.75 | 1.2817e-08
[5 empty] | 49.9 | 49.85 | 9.4757e-09
[5 empty] | 50 | 49.95 | 5.8769e-09
[5 empty] | 50.1 | 50.05 | 2.0185e-09

## MC, water, 1 cm, Hiatt
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
22 April 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Axxent
Manufacturer: | Xoft
Spectra information
origin | Monte Carlo simulation
Monte Carlo (MC) code | MCNP5 version 1.6
MC simulation settings | MCPLIB84 photon library, coupled photon-electron transport was required, and the bremsstrahlung weighting was increased by a factor of 50 (i.e., bnum and xnum cards) with high-energy preference for the source anode and its tungsten-film coating, i.e., bbrem b49 = 100). The default 1 keV photon and electron cut-off energies were used. The ITS-style indexing on the DBCN card was employed
authors | J. R. Hiatt et al.
location | see publication
publication | A revised dosimetric characterization of the model S700 electronic brachytherapy source containing an anode-centering plastic insert and other components not included in the 2006 model
DOI | http://dx.doi.org/10.1118/1.4919280
year of the result | 2015
material | water
source-to-detector distance (cm) | 1* |  | * see supplementary material of the publication for other distances and angles
description of source-to-detector distance | laterally from beam axis at the centre of anode cone* (marked by cross in the visualization of the geometry)
applicator | none
other information | the device is inserted inside a water-cooling catheter
Device setting
nominal bias voltage (kV) | 50
bias voltage sigma (kV) | N/A
anode current (mA) | 0.3
Visualization of the geometry
Spectra parameters
Filtration material and density (g/cm3) | Al |  | 2.7
Emean (in F; keV) |  |  | 28.815 | average photon energy
Emean (in Kair; keV) |  |  | 23.7937 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 21.701 | energy of monoenergetic photons with the same HVL as the given spectrum
1st HVL in Al (in Kair; mm) |  |  | 0.94381 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value
applicator diameter (mm) |  |  | none
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative)
4.95
| 5.05 | 5 | 0
| 5.15 | 5.1 | 1.706e-11
| 5.25 | 5.2 | 1.2792e-11
| 5.35 | 5.3 | 6.5291e-12
| 5.45 | 5.4 | 0
| 5.55 | 5.5 | 1.1341e-11
| 5.65 | 5.6 | 0
| 5.75 | 5.7 | 0
| 5.85 | 5.8 | 2.2194e-11
| 5.95 | 5.9 | 0
| 6.05 | 6 | 0
| 6.15 | 6.1 | 0
| 6.25 | 6.2 | 2.317e-12
| 6.35 | 6.3 | 1.0179e-11
| 6.45 | 6.4 | 2.9706e-12
| 6.55 | 6.5 | 1.9626e-11
| 6.65 | 6.6 | 2.8542e-12
| 6.75 | 6.7 | 0
| 6.85 | 6.8 | 1.189e-11
| 6.95 | 6.9 | 0
| 7.05 | 7 | 0
| 7.15 | 7.1 | 0
| 7.25 | 7.2 | 3.8085e-11
| 7.35 | 7.3 | 0
| 7.45 | 7.4 | 4.3338e-11
| 7.55 | 7.5 | 0
| 7.65 | 7.6 | 9.5261e-12
| 7.75 | 7.7 | 1.247e-11
| 7.85 | 7.8 | 3.236e-12
| 7.95 | 7.9 | 2.1201e-11
| 8.05 | 8 | 3.0166e-12
| 8.15 | 8.1 | 1.6133e-11
| 8.25 | 8.2 | 2.9245e-11
| 8.35 | 8.3 | 1.0137e-11
| 8.45 | 8.4 | 4.3204e-11
| 8.55 | 8.5 | 1.9269e-11
| 8.65 | 8.6 | 4.1684e-11
| 8.75 | 8.7 | 7.0549e-11
| 8.85 | 8.8 | 6.3921e-10
| 8.95 | 8.9 | 1.1714e-10
| 9.05 | 9 | 2.228e-10
| 9.15 | 9.1 | 2.2504e-10
| 9.25 | 9.2 | 3.3018e-10
| 9.35 | 9.3 | 4.8388e-10
| 9.45 | 9.4 | 6.043e-10
| 9.55 | 9.5 | 9.2214e-10
| 9.65 | 9.6 | 9.5163e-10
| 9.75 | 9.7 | 1.1739e-09
| 9.85 | 9.8 | 1.6362e-09
| 9.95 | 9.9 | 2.0151e-09
| 10.05 | 10 | 2.2754e-09
| 10.15 | 10.1 | 2.8208e-09
| 10.25 | 10.2 | 3.2952e-09
| 10.35 | 10.3 | 2.9066e-09
| 10.45 | 10.4 | 3.5553e-09
| 10.55 | 10.5 | 4.3822e-09
| 10.65 | 10.6 | 5.316e-09
| 10.75 | 10.7 | 6.6572e-09
| 10.85 | 10.8 | 7.1684e-09
| 10.95 | 10.9 | 8.1261e-09
| 11.05 | 11 | 9.6763e-09
| 11.15 | 11.1 | 1.1551e-08
| 11.25 | 11.2 | 1.3249e-08
| 11.35 | 11.3 | 1.5199e-08
| 11.45 | 11.4 | 1.728e-08
| 11.55 | 11.5 | 1.9225e-08
| 11.65 | 11.6 | 1.9371e-08
| 11.75 | 11.7 | 2.1332e-08
| 11.85 | 11.8 | 2.4134e-08
| 11.95 | 11.9 | 2.6968e-08
| 12.05 | 12 | 3.0024e-08
| 12.15 | 12.1 | 3.1285e-08
| 12.25 | 12.2 | 3.4373e-08
| 12.35 | 12.3 | 3.7298e-08
| 12.45 | 12.4 | 4.137e-08
| 12.55 | 12.5 | 4.5599e-08
| 12.65 | 12.6 | 5.0077e-08
| 12.75 | 12.7 | 5.3328e-08
| 12.85 | 12.8 | 5.7577e-08
| 12.95 | 12.9 | 6.2563e-08
| 13.05 | 13 | 6.8306e-08
| 13.15 | 13.1 | 7.3016e-08
| 13.25 | 13.2 | 7.8237e-08
| 13.35 | 13.3 | 8.3707e-08
| 13.45 | 13.4 | 9.0046e-08
| 13.55 | 13.5 | 9.7635e-08
| 13.65 | 13.6 | 1.037e-07
| 13.75 | 13.7 | 1.1003e-07
| 13.85 | 13.8 | 1.1878e-07
| 13.95 | 13.9 | 1.2606e-07
| 14.05 | 14 | 1.3522e-07
| 14.15 | 14.1 | 1.4894e-07
| 14.25 | 14.2 | 1.6277e-07
| 14.35 | 14.3 | 1.7207e-07
| 14.45 | 14.4 | 1.7871e-07
| 14.55 | 14.5 | 1.91e-07
| 14.65 | 14.6 | 1.9859e-07
| 14.75 | 14.7 | 2.0843e-07
| 14.85 | 14.8 | 6.676e-07
| 14.95 | 14.9 | 1.1011e-06
| 15.05 | 15 | 2.0848e-07
| 15.15 | 15.1 | 2.1349e-07
| 15.25 | 15.2 | 2.2313e-07
| 15.35 | 15.3 | 2.3217e-07
| 15.45 | 15.4 | 2.4008e-07
| 15.55 | 15.5 | 2.5128e-07
| 15.65 | 15.6 | 2.5954e-07
| 15.75 | 15.7 | 2.7171e-07
| 15.85 | 15.8 | 2.8149e-07
| 15.95 | 15.9 | 2.857e-07
| 16.05 | 16 | 2.9878e-07
| 16.15 | 16.1 | 3.05e-07
| 16.25 | 16.2 | 3.107e-07
| 16.35 | 16.3 | 3.1854e-07
| 16.45 | 16.4 | 3.2494e-07
| 16.55 | 16.5 | 3.2844e-07
| 16.65 | 16.6 | 3.3666e-07
| 16.75 | 16.7 | 7.1192e-07
| 16.85 | 16.8 | 3.4383e-07
| 16.95 | 16.9 | 3.5299e-07
| 17.05 | 17 | 3.5186e-07
| 17.15 | 17.1 | 2.3634e-07
| 17.25 | 17.2 | 2.4381e-07
| 17.35 | 17.3 | 2.5158e-07
| 17.45 | 17.4 | 2.5742e-07
| 17.55 | 17.5 | 2.6573e-07
| 17.65 | 17.6 | 2.7255e-07
| 17.75 | 17.7 | 2.7926e-07
| 17.85 | 17.8 | 2.8411e-07
| 17.95 | 17.9 | 2.9259e-07
| 18.05 | 18 | 2.9888e-07
| 18.15 | 18.1 | 3.0369e-07
| 18.25 | 18.2 | 3.1014e-07
| 18.35 | 18.3 | 3.1842e-07
| 18.45 | 18.4 | 3.2773e-07
| 18.55 | 18.5 | 3.3362e-07
| 18.65 | 18.6 | 3.3995e-07
| 18.75 | 18.7 | 3.4413e-07
| 18.85 | 18.8 | 3.5377e-07
| 18.95 | 18.9 | 3.6012e-07
| 19.05 | 19 | 3.6717e-07
| 19.15 | 19.1 | 3.7196e-07
| 19.25 | 19.2 | 3.7899e-07
| 19.35 | 19.3 | 3.8694e-07
| 19.45 | 19.4 | 3.9123e-07
| 19.55 | 19.5 | 3.9973e-07
| 19.65 | 19.6 | 4.0207e-07
| 19.75 | 19.7 | 4.1175e-07
| 19.85 | 19.8 | 4.1701e-07
| 19.95 | 19.9 | 4.2483e-07
| 20.05 | 20 | 4.3222e-07
| 20.15 | 20.1 | 4.3472e-07
| 20.25 | 20.2 | 4.4411e-07
| 20.35 | 20.3 | 4.4809e-07
| 20.45 | 20.4 | 4.5875e-07
| 20.55 | 20.5 | 4.5737e-07
| 20.65 | 20.6 | 4.6354e-07
| 20.75 | 20.7 | 4.6962e-07
| 20.85 | 20.8 | 4.7415e-07
| 20.95 | 20.9 | 4.8341e-07
| 21.05 | 21 | 4.8299e-07
| 21.15 | 21.1 | 4.8771e-07
| 21.25 | 21.2 | 4.9409e-07
| 21.35 | 21.3 | 4.9993e-07
| 21.45 | 21.4 | 5.0372e-07
| 21.55 | 21.5 | 5.0646e-07
| 21.65 | 21.6 | 5.1386e-07
| 21.75 | 21.7 | 5.2028e-07
| 21.85 | 21.8 | 5.2027e-07
| 21.95 | 21.9 | 5.2592e-07
| 22.05 | 22 | 6.1274e-07
| 22.15 | 22.1 | 6.9168e-07
| 22.25 | 22.2 | 5.3415e-07
| 22.35 | 22.3 | 5.4092e-07
| 22.45 | 22.4 | 5.4214e-07
| 22.55 | 22.5 | 5.4495e-07
| 22.65 | 22.6 | 5.4973e-07
| 22.75 | 22.7 | 5.5308e-07
| 22.85 | 22.8 | 5.5645e-07
| 22.95 | 22.9 | 5.5782e-07
| 23.05 | 23 | 5.5914e-07
| 23.15 | 23.1 | 5.674e-07
| 23.25 | 23.2 | 5.6554e-07
| 23.35 | 23.3 | 5.7307e-07
| 23.45 | 23.4 | 5.7321e-07
| 23.55 | 23.5 | 5.7475e-07
| 23.65 | 23.6 | 5.7571e-07
| 23.75 | 23.7 | 5.8469e-07
| 23.85 | 23.8 | 5.8186e-07
| 23.95 | 23.9 | 5.8199e-07
| 24.05 | 24 | 5.8514e-07
| 24.15 | 24.1 | 5.8975e-07
| 24.25 | 24.2 | 5.8915e-07
| 24.35 | 24.3 | 5.9331e-07
| 24.45 | 24.4 | 5.9374e-07
| 24.55 | 24.5 | 5.9387e-07
| 24.65 | 24.6 | 6.0079e-07
| 24.75 | 24.7 | 5.9743e-07
| 24.85 | 24.8 | 6.044e-07
| 24.95 | 24.9 | 6.5913e-07
| 25.05 | 25 | 5.9963e-07
| 25.15 | 25.1 | 6.0196e-07
| 25.25 | 25.2 | 6.008e-07
| 25.35 | 25.3 | 6.0469e-07
| 25.45 | 25.4 | 6.193e-07
| 25.55 | 25.5 | 6.0611e-07
| 25.65 | 25.6 | 6.0592e-07
| 25.75 | 25.7 | 6.1067e-07
| 25.85 | 25.8 | 6.0697e-07
| 25.95 | 25.9 | 6.0786e-07
| 26.05 | 26 | 6.1161e-07
| 26.15 | 26.1 | 6.0691e-07
| 26.25 | 26.2 | 6.0716e-07
| 26.35 | 26.3 | 6.1086e-07
| 26.45 | 26.4 | 6.1052e-07
| 26.55 | 26.5 | 6.1256e-07
| 26.65 | 26.6 | 6.1394e-07
| 26.75 | 26.7 | 6.1073e-07
| 26.85 | 26.8 | 6.0922e-07
| 26.95 | 26.9 | 6.1103e-07
| 27.05 | 27 | 6.1101e-07
| 27.15 | 27.1 | 6.1037e-07
| 27.25 | 27.2 | 6.0888e-07
| 27.35 | 27.3 | 6.091e-07
| 27.45 | 27.4 | 6.0932e-07
| 27.55 | 27.5 | 6.0719e-07
| 27.65 | 27.6 | 6.0989e-07
| 27.75 | 27.7 | 6.0961e-07
| 27.85 | 27.8 | 6.0801e-07
| 27.95 | 27.9 | 6.0705e-07
| 28.05 | 28 | 6.0614e-07
| 28.15 | 28.1 | 6.0461e-07
| 28.25 | 28.2 | 6.0471e-07
| 28.35 | 28.3 | 6.0109e-07
| 28.45 | 28.4 | 6.0214e-07
| 28.55 | 28.5 | 6.0032e-07
| 28.65 | 28.6 | 6.0048e-07
| 28.75 | 28.7 | 6.0012e-07
| 28.85 | 28.8 | 5.9699e-07
| 28.95 | 28.9 | 5.9834e-07
| 29.05 | 29 | 5.9668e-07
| 29.15 | 29.1 | 5.9368e-07
| 29.25 | 29.2 | 5.9351e-07
| 29.35 | 29.3 | 5.8982e-07
| 29.45 | 29.4 | 5.8846e-07
| 29.55 | 29.5 | 5.8894e-07
| 29.65 | 29.6 | 5.8723e-07
| 29.75 | 29.7 | 5.8497e-07
| 29.85 | 29.8 | 5.8182e-07
| 29.95 | 29.9 | 5.8018e-07
| 30.05 | 30 | 5.8168e-07
| 30.15 | 30.1 | 5.7424e-07
| 30.25 | 30.2 | 5.7849e-07
| 30.35 | 30.3 | 5.7449e-07
| 30.45 | 30.4 | 5.7106e-07
| 30.55 | 30.5 | 5.7211e-07
| 30.65 | 30.6 | 5.7059e-07
| 30.75 | 30.7 | 5.6372e-07
| 30.85 | 30.8 | 5.6392e-07
| 30.95 | 30.9 | 5.5973e-07
| 31.05 | 31 | 5.5895e-07
| 31.15 | 31.1 | 5.5891e-07
| 31.25 | 31.2 | 5.5284e-07
| 31.35 | 31.3 | 5.5194e-07
| 31.45 | 31.4 | 5.4911e-07
| 31.55 | 31.5 | 5.4931e-07
| 31.65 | 31.6 | 5.4578e-07
| 31.75 | 31.7 | 5.4511e-07
| 31.85 | 31.8 | 5.4203e-07
| 31.95 | 31.9 | 5.4202e-07
| 32.05 | 32 | 5.3437e-07
| 32.15 | 32.1 | 5.3251e-07
| 32.25 | 32.2 | 5.3285e-07
| 32.35 | 32.3 | 5.2732e-07
| 32.45 | 32.4 | 5.2339e-07
| 32.55 | 32.5 | 5.2172e-07
| 32.65 | 32.6 | 5.1802e-07
| 32.75 | 32.7 | 5.1943e-07
| 32.85 | 32.8 | 5.1518e-07
| 32.95 | 32.9 | 5.1398e-07
| 33.05 | 33 | 5.1116e-07
| 33.15 | 33.1 | 5.0767e-07
| 33.25 | 33.2 | 5.0394e-07
| 33.35 | 33.3 | 5.0065e-07
| 33.45 | 33.4 | 5.0013e-07
| 33.55 | 33.5 | 4.9663e-07
| 33.65 | 33.6 | 4.9401e-07
| 33.75 | 33.7 | 4.9188e-07
| 33.85 | 33.8 | 4.8877e-07
| 33.95 | 33.9 | 4.8569e-07
| 34.05 | 34 | 4.8385e-07
| 34.15 | 34.1 | 4.814e-07
| 34.25 | 34.2 | 4.7497e-07
| 34.35 | 34.3 | 4.7364e-07
| 34.45 | 34.4 | 4.7377e-07
| 34.55 | 34.5 | 4.6684e-07
| 34.65 | 34.6 | 4.6379e-07
| 34.75 | 34.7 | 4.6256e-07
| 34.85 | 34.8 | 4.5785e-07
| 34.95 | 34.9 | 4.5565e-07
| 35.05 | 35 | 4.5174e-07
| 35.15 | 35.1 | 4.4693e-07
| 35.25 | 35.2 | 4.456e-07
| 35.35 | 35.3 | 4.4179e-07
| 35.45 | 35.4 | 4.4018e-07
| 35.55 | 35.5 | 4.3556e-07
| 35.65 | 35.6 | 4.3221e-07
| 35.75 | 35.7 | 4.2943e-07
| 35.85 | 35.8 | 4.2552e-07
| 35.95 | 35.9 | 4.231e-07
| 36.05 | 36 | 4.2257e-07
| 36.15 | 36.1 | 4.1971e-07
| 36.25 | 36.2 | 4.1554e-07
| 36.35 | 36.3 | 4.125e-07
| 36.45 | 36.4 | 4.0965e-07
| 36.55 | 36.5 | 4.0549e-07
| 36.65 | 36.6 | 4.0096e-07
| 36.75 | 36.7 | 3.9786e-07
| 36.85 | 36.8 | 3.949e-07
| 36.95 | 36.9 | 3.9193e-07
| 37.05 | 37 | 3.895e-07
| 37.15 | 37.1 | 3.8504e-07
| 37.25 | 37.2 | 3.8355e-07
| 37.35 | 37.3 | 3.7728e-07
| 37.45 | 37.4 | 3.7656e-07
| 37.55 | 37.5 | 3.7252e-07
| 37.65 | 37.6 | 3.6908e-07
| 37.75 | 37.7 | 3.6699e-07
| 37.85 | 37.8 | 3.6244e-07
| 37.95 | 37.9 | 3.6011e-07
| 38.05 | 38 | 3.5704e-07
| 38.15 | 38.1 | 3.5159e-07
| 38.25 | 38.2 | 3.4903e-07
| 38.35 | 38.3 | 3.455e-07
| 38.45 | 38.4 | 3.425e-07
| 38.55 | 38.5 | 3.4027e-07
| 38.65 | 38.6 | 3.3742e-07
| 38.75 | 38.7 | 3.3256e-07
| 38.85 | 38.8 | 3.2898e-07
| 38.95 | 38.9 | 3.2701e-07
| 39.05 | 39 | 3.2351e-07
| 39.15 | 39.1 | 3.1936e-07
| 39.25 | 39.2 | 3.1531e-07
| 39.35 | 39.3 | 3.139e-07
| 39.45 | 39.4 | 3.1067e-07
| 39.55 | 39.5 | 3.0885e-07
| 39.65 | 39.6 | 3.0275e-07
| 39.75 | 39.7 | 3.0102e-07
| 39.85 | 39.8 | 2.96e-07
| 39.95 | 39.9 | 2.9272e-07
| 40.05 | 40 | 2.8945e-07
| 40.15 | 40.1 | 2.8681e-07
| 40.25 | 40.2 | 2.8415e-07
| 40.35 | 40.3 | 2.8046e-07
| 40.45 | 40.4 | 2.7796e-07
| 40.55 | 40.5 | 2.7437e-07
| 40.65 | 40.6 | 2.7278e-07
| 40.75 | 40.7 | 2.6841e-07
| 40.85 | 40.8 | 2.6274e-07
| 40.95 | 40.9 | 2.6131e-07
| 41.05 | 41 | 2.5692e-07
| 41.15 | 41.1 | 2.5688e-07
| 41.25 | 41.2 | 2.5101e-07
| 41.35 | 41.3 | 2.4778e-07
| 41.45 | 41.4 | 2.4492e-07
| 41.55 | 41.5 | 2.4209e-07
| 41.65 | 41.6 | 2.3827e-07
| 41.75 | 41.7 | 2.342e-07
| 41.85 | 41.8 | 2.3088e-07
| 41.95 | 41.9 | 2.2816e-07
| 42.05 | 42 | 2.2568e-07
| 42.15 | 42.1 | 2.2048e-07
| 42.25 | 42.2 | 2.1786e-07
| 42.35 | 42.3 | 2.1556e-07
| 42.45 | 42.4 | 2.1105e-07
| 42.55 | 42.5 | 2.0849e-07
| 42.65 | 42.6 | 2.0471e-07
| 42.75 | 42.7 | 2.0245e-07
| 42.85 | 42.8 | 1.9998e-07
| 42.95 | 42.9 | 1.973e-07
| 43.05 | 43 | 1.9457e-07
| 43.15 | 43.1 | 1.9087e-07
| 43.25 | 43.2 | 1.8764e-07
| 43.35 | 43.3 | 1.8517e-07
| 43.45 | 43.4 | 1.8284e-07
| 43.55 | 43.5 | 1.7884e-07
| 43.65 | 43.6 | 1.7713e-07
| 43.75 | 43.7 | 1.7516e-07
| 43.85 | 43.8 | 1.7106e-07
| 43.95 | 43.9 | 1.6903e-07
| 44.05 | 44 | 1.6595e-07
| 44.15 | 44.1 | 1.6391e-07
| 44.25 | 44.2 | 1.613e-07
| 44.35 | 44.3 | 1.5745e-07
| 44.45 | 44.4 | 1.5555e-07
| 44.55 | 44.5 | 1.5346e-07
| 44.65 | 44.6 | 1.506e-07
| 44.75 | 44.7 | 1.482e-07
| 44.85 | 44.8 | 1.4608e-07
| 44.95 | 44.9 | 1.4331e-07
| 45.05 | 45 | 1.3997e-07
| 45.15 | 45.1 | 1.3706e-07
| 45.25 | 45.2 | 1.3363e-07
| 45.35 | 45.3 | 1.3105e-07
| 45.45 | 45.4 | 1.2897e-07
| 45.55 | 45.5 | 1.2642e-07
| 45.65 | 45.6 | 1.2397e-07
| 45.75 | 45.7 | 1.2129e-07
| 45.85 | 45.8 | 1.1829e-07
| 45.95 | 45.9 | 1.1716e-07
| 46.05 | 46 | 1.1409e-07
| 46.15 | 46.1 | 1.1123e-07
| 46.25 | 46.2 | 1.0949e-07
| 46.35 | 46.3 | 1.0738e-07
| 46.45 | 46.4 | 1.0403e-07
| 46.55 | 46.5 | 1.0172e-07
| 46.65 | 46.6 | 9.9743e-08
| 46.75 | 46.7 | 9.6316e-08
| 46.85 | 46.8 | 9.391e-08
| 46.95 | 46.9 | 9.1818e-08
| 47.05 | 47 | 8.9434e-08
| 47.15 | 47.1 | 8.6095e-08
| 47.25 | 47.2 | 8.3211e-08
| 47.35 | 47.3 | 8.0651e-08
| 47.45 | 47.4 | 7.8421e-08
| 47.55 | 47.5 | 7.5997e-08
| 47.65 | 47.6 | 7.3419e-08
| 47.75 | 47.7 | 7.0366e-08
| 47.85 | 47.8 | 6.8589e-08
| 47.95 | 47.9 | 6.5842e-08
| 48.05 | 48 | 6.3261e-08
| 48.15 | 48.1 | 6.1282e-08
| 48.25 | 48.2 | 5.8011e-08
| 48.35 | 48.3 | 5.3927e-08
| 48.45 | 48.4 | 5.2594e-08
| 48.55 | 48.5 | 4.9448e-08
| 48.65 | 48.6 | 4.641e-08
| 48.75 | 48.7 | 4.4437e-08
| 48.85 | 48.8 | 4.1458e-08
| 48.95 | 48.9 | 3.8459e-08
| 49.05 | 49 | 3.5487e-08
| 49.15 | 49.1 | 3.2789e-08
| 49.25 | 49.2 | 3.0113e-08
| 49.35 | 49.3 | 2.7328e-08
| 49.45 | 49.4 | 2.4202e-08
| 49.55 | 49.5 | 2.1082e-08
| 49.65 | 49.6 | 1.7517e-08
| 49.75 | 49.7 | 1.431e-08
| 49.85 | 49.8 | 1.1029e-08
| 49.95 | 49.9 | 7.1707e-09
| 50.05 | 50 | 2.3043e-09
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Graphical only
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
22 April 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Axxent | spectra available in graphical form only
Manufacturer: | Xoft
SPECTRA SET 1 | Spectra information
origin | measurement
detector | Canberra liquid-nitrogen-cooled Ultra Low-Energy HPGe type GUL0110. The detector incorporates an HPGe crystal with an active diameter of 11.3 mm and a thickness of 10 mm, behind a 0.025 mm Be cryostat window
data evaluation | a simple backward-stripping algorithm was used to obtain fluence spectrum from the detector spectrum
authors | NIST
location | USA
publication | New National Air-Kerma Standard for Low-Energy Electronic Brachytherapy Sources
DOI | http://dx.doi.org/10.6028/jres.119.022
year of the result | 2014
material | air
source-to-detector distance (cm) | 50
description of source-to-detector distance | laterally from beam axis at the tip depth
applicator | none
other information | the device is inserted inside a water-cooling catheter
Device setting
nominal bias voltage (kV) | 50
bias voltage sigma (kV) | N/A
anode current (mA) | 0.3