## Description
General information
Catalogue version: 6 October 2021
Created by Jaroslav Šolc, jsolc@cmi.cz, Czech Metrology Institute
This Excel file contains the catalogue of X-ray energy spectra of electronic brachytherapy (eBT) device Esteya manufactured by Electa.
The catalogue was compiled from suitable data available in literature and from the data measured or created by the project "18NRM02 PRISM-eBT" (Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices; http://www.ebt-empir.eu/) of the European Metrology Programme for Innovation and Research (EMPIR).
The main goal of the catalogue was to develop a set of harmonised reference X-ray qualities denoted as “eBT equivalent” that would be established for in-laboratory activities by the X-ray tubes usually available for calibration in calibration laboratories, but the tube voltage and the filter are specifically selected to simulate the energy spectrum of the corresponding eBT radiation as closely as possible at the reference point.
In general, the catalogue consists of:
1) X-ray photon fluence eBT spectra measured and/or Monte Carlo simulated in air and/or in water, whenever they were available;
2) eBT-equivalent X-ray photon fluence spectra that match selected eBT spectra. Typically, the eBT-equivalent spectra were calculated by a dedicated software, usually SpekPy software. If available, the eBT-equivalent spectra already realized are also included. Please note that the eBT-equivalent spectra may be site-dependent, based on properties of X-ray generator available (anode angle, inherent filtration) and the distance from anode to measurement point (attenuation in air). Realized eBT-equivalent spectra should be validated by spectrometry.
Specifically, this catalog includes spectra listed below:
Sheet | Description
EXP, air, 100 cm, 70 kV | Measured spectrum in air, 100 cm source-to-detector distance
EXP, air, 150 cm, 70 kV, VSL | 1) Measured spectrum in air, 150 cm source-to-detector distance 2) eBT-equivalent spectrum in air, 150 cm source-to-detector distance, calculated by SpekCalc software 3) Intermediate results for eBT-equivalent spectrum in air, 150 cm source-to-detector distance, realized at VSL

## EXP, air, 100 cm, 70 kV
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
5 May 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Esteya
Manufacturer: | Elekta
Spectra information
origin | measurement
detector | Ortec GLP 25300/13P4, Ge crystal: diameter 25 mm, thickness 13 mm
data evaluation | none, raw measured detector spectrum
authors | manufacturer's data, Yury Niatsetski, private communication
location | -
publication | also available in: Dosimetric characteristics of a new unit for electronic skin brachytherapy
DOI | https://doi.org/10.5114/jcb.2014.40770
year of the result | 2014
material | air
source-to-detector distance (cm) | 100 | (approximately)
description of source-to-detector distance | on device axis
applicator | 30 mm
other information | measurement behind pinholes, average over three different applicators
Device setting
nominal bias voltage (kV) | 70
bias voltage sigma (kV) | N/A
anode current (mA) | N/A
Spectra parameters
Filtration material and density (g/cm3) | Al |  | 2.7
Emean (in F; keV) |  |  | 36.1066 | average photon energy
Emean (in Kair; keV) |  |  | 29.7991 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 26.7325 | energy of monoenergetic photons with the same HVL as the given spectrum | 99999.9998
1st HVL in Al (in Kair; mm) |  |  | 1.6813 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value | 58.2174
applicator diameter (mm) |  |  | 30
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative)
14
| 14.5 | 14.25 | 89.1422 |  | 0.051896
| 15 | 14.75 | 124.3987 |  | 0.072422
| 15.5 | 15.25 | 168.797 |  | 0.098269
| 16 | 15.75 | 223.3393 |  | 0.13002
| 16.5 | 16.25 | 291.9763 |  | 0.16998
| 17 | 16.75 | 361.4559 |  | 0.21043
| 17.5 | 17.25 | 441.6996 |  | 0.25715
| 18 | 17.75 | 528.4881 |  | 0.30767
| 18.5 | 18.25 | 619.749 |  | 0.3608
| 19 | 18.75 | 708.0684 |  | 0.41222
| 19.5 | 19.25 | 811.6881 |  | 0.47254
| 20 | 19.75 | 904.4018 |  | 0.52652
| 20.5 | 20.25 | 994.5955 |  | 0.57903
| 21 | 20.75 | 1074.0995 |  | 0.62531
| 21.5 | 21.25 | 1153.261 |  | 0.6714
| 22 | 21.75 | 1232.6099 |  | 0.71759
| 22.5 | 22.25 | 1295.689 |  | 0.75432
| 23 | 22.75 | 1362.2212 |  | 0.79305
| 23.5 | 23.25 | 1420.2537 |  | 0.82684
| 24 | 23.75 | 1475.5342 |  | 0.85902
| 24.5 | 24.25 | 1510.2228 |  | 0.87921
| 25 | 24.75 | 1562.3358 |  | 0.90955
| 25.5 | 25.25 | 1582.7645 |  | 0.92144
| 26 | 25.75 | 1623.1576 |  | 0.94496
| 26.5 | 26.25 | 1639.6914 |  | 0.95459
| 27 | 26.75 | 1657.726 |  | 0.96509
| 27.5 | 27.25 | 1679.5419 |  | 0.97779
| 28 | 27.75 | 1696.9583 |  | 0.98793
| 28.5 | 28.25 | 1702.3117 |  | 0.99104
| 29 | 28.75 | 1698.4003 |  | 0.98876
| 29.5 | 29.25 | 1701.2963 |  | 0.99045
| 30 | 29.75 | 1702.2569 |  | 0.99101
| 30.5 | 30.25 | 1690.5078 |  | 0.98417
| 31 | 30.75 | 1683.4018 |  | 0.98003
| 31.5 | 31.25 | 1681.7246 |  | 0.97906
| 32 | 31.75 | 1666.6228 |  | 0.97026
| 32.5 | 32.25 | 1667.7067 |  | 0.9709
| 33 | 32.75 | 1635.2306 |  | 0.95199
| 33.5 | 33.25 | 1617.977 |  | 0.94194
| 34 | 33.75 | 1590.3857 |  | 0.92588
| 34.5 | 34.25 | 1566.5236 |  | 0.91199
| 35 | 34.75 | 1554.8788 |  | 0.90521
| 35.5 | 35.25 | 1521.3883 |  | 0.88571
| 36 | 35.75 | 1511.9942 |  | 0.88024
| 36.5 | 36.25 | 1487.4749 |  | 0.86597
| 37 | 36.75 | 1465.3423 |  | 0.85308
| 37.5 | 37.25 | 1432.1643 |  | 0.83377
| 38 | 37.75 | 1403.9056 |  | 0.81732
| 38.5 | 38.25 | 1383.291 |  | 0.80532
| 39 | 38.75 | 1352.8646 |  | 0.7876
| 39.5 | 39.25 | 1329.9749 |  | 0.77428
| 40 | 39.75 | 1298.9261 |  | 0.7562
| 40.5 | 40.25 | 1272.1174 |  | 0.74059
| 41 | 40.75 | 1245.6532 |  | 0.72519
| 41.5 | 41.25 | 1215.4782 |  | 0.70762
| 42 | 41.75 | 1195.0778 |  | 0.69574
| 42.5 | 42.25 | 1168.2207 |  | 0.68011
| 43 | 42.75 | 1141.9278 |  | 0.6648
| 43.5 | 43.25 | 1111.8373 |  | 0.64728
| 44 | 43.75 | 1084.0836 |  | 0.63113
| 44.5 | 44.25 | 1064.078 |  | 0.61948
| 45 | 44.75 | 1033.8127 |  | 0.60186
| 45.5 | 45.25 | 1003.1517 |  | 0.58401
| 46 | 45.75 | 979.9465 |  | 0.5705
| 46.5 | 46.25 | 957.9569 |  | 0.5577
| 47 | 46.75 | 925.0789 |  | 0.53856
| 47.5 | 47.25 | 909.7017 |  | 0.5296
| 48 | 47.75 | 881.5007 |  | 0.51319
| 48.5 | 48.25 | 856.3677 |  | 0.49856
| 49 | 48.75 | 829.1017 |  | 0.48268
| 49.5 | 49.25 | 804.1912 |  | 0.46818
| 50 | 49.75 | 781.7221 |  | 0.4551
| 50.5 | 50.25 | 759.2392 |  | 0.44201
| 51 | 50.75 | 740.8752 |  | 0.43132
| 51.5 | 51.25 | 711.2858 |  | 0.41409
| 52 | 51.75 | 686.0768 |  | 0.39942
| 52.5 | 52.25 | 664.259 |  | 0.38671
| 53 | 52.75 | 640.0527 |  | 0.37262
| 53.5 | 53.25 | 617.2226 |  | 0.35933
| 54 | 53.75 | 593.4761 |  | 0.34551
| 54.5 | 54.25 | 574.4643 |  | 0.33444
| 55 | 54.75 | 549.7926 |  | 0.32008
| 55.5 | 55.25 | 530.383 |  | 0.30878
| 56 | 55.75 | 510.3462 |  | 0.29711
| 56.5 | 56.25 | 486.2165 |  | 0.28306
| 57 | 56.75 | 467.4811 |  | 0.27216
| 57.5 | 57.25 | 445.0178 |  | 0.25908
| 58 | 57.75 | 424.7771 |  | 0.24729
| 58.5 | 58.25 | 402.7386 |  | 0.23446
| 59 | 58.75 | 380.9744 |  | 0.22179
| 59.5 | 59.25 | 363.0456 |  | 0.21136
| 60 | 59.75 | 346.8614 |  | 0.20193
| 60.5 | 60.25 | 327.664 |  | 0.19076
| 61 | 60.75 | 304.3561 |  | 0.17719
| 61.5 | 61.25 | 283.3285 |  | 0.16495
| 62 | 61.75 | 264.1689 |  | 0.15379
| 62.5 | 62.25 | 247.8608 |  | 0.1443
| 63 | 62.75 | 229.5341 |  | 0.13363
| 63.5 | 63.25 | 212.4485 |  | 0.12368
| 64 | 63.75 | 192.1299 |  | 0.11185
| 64.5 | 64.25 | 176.8141 |  | 0.10294
| 65 | 64.75 | 155.6473 |  | 0.090614
| 65.5 | 65.25 | 140.3068 |  | 0.081683
| 66 | 65.75 | 120.7032 |  | 0.07027
| 66.5 | 66.25 | 105.0657 |  | 0.061167
| 67 | 66.75 | 89.2202 |  | 0.051942
| 67.5 | 67.25 | 73.4117 |  | 0.042738
| 68 | 67.75 | 56.074 |  | 0.032645
| 68.5 | 68.25 | 42.1193 |  | 0.024521
| 69 | 68.75 | 27.0348 |  | 0.015739
| 69.5 | 69.25 | 16.1027 |  | 0.0093746
| 70 | 69.75 | 0 |  | 0

## EXP, air, 150 cm, 70 kV, VSL
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
12 July 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Esteya |  |  |  | This sheet also contains eBT-equivalent spectra
Manufacturer: | Elekta
Spectra information
origin | measurement
detector | Ortec HPGe detector
data evaluation | average of 10 measured detector spectra
authors | Leon de Prez
location | VSL
publication | N/A
DOI | N/A
year of the result | 2021
material | air
source-to-detector distance (cm) | 150
description of source-to-detector distance | on central axis, distance from source
applicator | 30 mm
other information | -
Device setting
nominal bias voltage (kV) | 70
bias voltage sigma (kV) | 1
current (mA) | N/A
eBT-equivalent spectra input parameters
Method | [6 empty] | SpekCalc fit |  |  | measurement
Anode material | [6 empty] | W |  |  | W
Anode angle (degree) | [6 empty] | 20 |  |  | 20
X-ray generator voltage (kWp) | [6 empty] | 70 |  |  | 70
Filtration 1 - material and thickness (mm) | Be | [5 empty] | 3 |  |  | 3
Filtration 2 - material and thickness (mm) | Air | [5 empty] | 1500 |  |  | 1500
Filtration 3 - material and thickness (mm) | Al | [5 empty] | 1.2 |  |  | 1.2
eBT and eBT-equivalent spectra characteristics
Filtration material, maximum thickness (mm; user-defined) and density (g/cm3) | Al | 2 | 2.7 |  |  |  | 2.7 |  |  | 2.7
Emean (in F; keV) |  |  | 36.0034 |  |  |  | 36.5108 |  |  | 36.2553 | average photon energy
Emean (in Kair; keV) |  |  | 29.6486 |  |  |  | 29.7181 |  |  | 29.261 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 26.5935 |  |  |  | 26.4562 |  |  | 25.9737 | energy of monoenergetic photons with the same HVL as the given spectrum | 29.1087 | 29.1087
1st HVL in Al - Excel (in Kair; mm) |  |  | 1.6581 |  |  |  | 1.6352 |  |  | 1.5562 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value; determined by Excel script | 30.4481 | 30.6541
[5 empty] | eBT-equivalent spectrum at 150 cm in air: SpekCalc intermediate results |  |  | eBT-equivalent spectrum at 150 cm in air: VSL realization intermediate results
applicator diameter (mm) |  |  | 30 |  |  |  | 30 |  |  | 30
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) |  | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative)
| 12.6338 |  |  |  | 9.5 |  |  | 11.7615
| 13.1338 | 12.8838 | 0.016676 |  | 10 | 9.75 | 4.8705e-05 | 12.2615 | 12.0115 | 0.022679 |  | 4.6563e-05 | 0.021535
| 13.6338 | 13.3838 | 0.028416 |  | 10.5 | 10.25 | 0.00015592 | 12.7615 | 12.5115 | 0.023497 |  | 0.00014906 | 0.022312
| 14.1338 | 13.8838 | 0.040559 |  | 11 | 10.75 | 0.00051467 | 13.2615 | 13.0115 | 0.032849 |  | 0.00049203 | 0.031193
| 14.6338 | 14.3838 | 0.055369 |  | 11.5 | 11.25 | 0.0012865 | 13.7615 | 13.5115 | 0.04863 |  | 0.0012299 | 0.046179
| 15.1338 | 14.8838 | 0.0766 |  | 12 | 11.75 | 0.003224 | 14.2615 | 14.0115 | 0.070569 |  | 0.0030822 | 0.067011
| 15.6338 | 15.3838 | 0.10561 |  | 12.5 | 12.25 | 0.006373 | 14.7615 | 14.5115 | 0.099365 |  | 0.0060926 | 0.094355
| 16.1338 | 15.8838 | 0.14183 |  | 13 | 12.75 | 0.012636 | 15.2615 | 15.0115 | 0.12633 |  | 0.01208 | 0.11996
| 16.6338 | 16.3838 | 0.17974 |  | 13.5 | 13.25 | 0.021611 | 15.7615 | 15.5115 | 0.16569 |  | 0.020661 | 0.15733
| 17.1338 | 16.8838 | 0.22246 |  | 14 | 13.75 | 0.037143 | 16.2615 | 16.0115 | 0.21081 |  | 0.035509 | 0.20018
| 17.6338 | 17.3838 | 0.27258 |  | 14.5 | 14.25 | 0.055341 | 16.7615 | 16.5115 | 0.25722 |  | 0.052907 | 0.24425
| 18.1338 | 17.8838 | 0.32646 |  | 15 | 14.75 | 0.082703 | 17.2615 | 17.0115 | 0.30063 |  | 0.079065 | 0.28547
| 18.6338 | 18.3838 | 0.37926 |  | 15.5 | 15.25 | 0.11236 | 17.7615 | 17.5115 | 0.35011 |  | 0.10742 | 0.33246
| 19.1338 | 18.8838 | 0.42831 |  | 16 | 15.75 | 0.15292 | 18.2615 | 18.0115 | 0.40521 |  | 0.14619 | 0.38479
| 19.6338 | 19.3838 | 0.48542 |  | 16.5 | 16.25 | 0.19257 | 18.7615 | 18.5115 | 0.45864 |  | 0.1841 | 0.43552
| 20.1338 | 19.8838 | 0.54189 |  | 17 | 16.75 | 0.24318 | 19.2615 | 19.0115 | 0.5142 |  | 0.23249 | 0.48828
| 20.6338 | 20.3838 | 0.58887 |  | 17.5 | 17.25 | 0.2898 | 19.7615 | 19.5115 | 0.56268 |  | 0.27705 | 0.53431
| 21.1338 | 20.8838 | 0.63895 |  | 18 | 17.75 | 0.34584 | 20.2615 | 20.0115 | 0.61053 |  | 0.33063 | 0.57975
| 21.6338 | 21.3838 | 0.68105 |  | 18.5 | 18.25 | 0.3966 | 20.7615 | 20.5115 | 0.6516 |  | 0.37915 | 0.61875
| 22.1338 | 21.8838 | 0.72663 |  | 19 | 18.75 | 0.45512 | 21.2615 | 21.0115 | 0.70046 |  | 0.4351 | 0.66515
| 22.6338 | 22.3838 | 0.76871 |  | 19.5 | 19.25 | 0.50588 | 21.7615 | 21.5115 | 0.74489 |  | 0.48363 | 0.70733
| 23.1338 | 22.8838 | 0.80747 |  | 20 | 19.75 | 0.5625 | 22.2615 | 22.0115 | 0.77458 |  | 0.53775 | 0.73553
| 23.6338 | 23.3838 | 0.83426 |  | 20.5 | 20.25 | 0.60995 | 22.7615 | 22.5115 | 0.81958 |  | 0.58312 | 0.77827
| 24.1338 | 23.8838 | 0.87035 |  | 21 | 20.75 | 0.66223 | 23.2615 | 23.0115 | 0.85707 |  | 0.63309 | 0.81386
| 24.6338 | 24.3838 | 0.90278 |  | 21.5 | 21.25 | 0.7044 | 23.7615 | 23.5115 | 0.88337 |  | 0.67341 | 0.83884
| 25.1338 | 24.8838 | 0.91277 |  | 22 | 21.75 | 0.74973 | 24.2615 | 24.0115 | 0.89113 |  | 0.71675 | 0.8462
| 25.6338 | 25.3838 | 0.93331 |  | 22.5 | 22.25 | 0.78632 | 24.7615 | 24.5115 | 0.92659 |  | 0.75173 | 0.87987
| 26.1338 | 25.8838 | 0.94453 |  | 23 | 22.75 | 0.82465 | 25.2615 | 25.0115 | 0.94334 |  | 0.78837 | 0.89578
| 26.6338 | 26.3838 | 0.95993 |  | 23.5 | 23.25 | 0.8538 | 25.7615 | 25.5115 | 0.96707 |  | 0.81624 | 0.91832
| 27.1338 | 26.8838 | 0.97577 |  | 24 | 23.75 | 0.88503 | 26.2615 | 26.0115 | 0.96166 |  | 0.8461 | 0.91318
| 27.6338 | 27.3838 | 0.98361 |  | 24.5 | 24.25 | 0.90591 | 26.7615 | 26.5115 | 0.98348 |  | 0.86606 | 0.9339
| 28.1338 | 27.8838 | 0.98601 |  | 25 | 24.75 | 0.9277 | 27.2615 | 27.0115 | 0.98352 |  | 0.88689 | 0.93393
| 28.6338 | 28.3838 | 0.98721 |  | 25.5 | 25.25 | 0.94476 | 27.7615 | 27.5115 | 0.99755 |  | 0.9032 | 0.94725
| 29.1338 | 28.8838 | 0.98963 |  | 26 | 25.75 | 0.96217 | 28.2615 | 28.0115 | 1 |  | 0.91984 | 0.94959
| 29.6338 | 29.3838 | 0.99019 |  | 26.5 | 26.25 | 0.97339 | 28.7615 | 28.5115 | 0.99451 |  | 0.93057 | 0.94437
| 30.1338 | 29.8838 | 0.99332 |  | 27 | 26.75 | 0.98455 | 29.2615 | 29.0115 | 0.99735 |  | 0.94124 | 0.94707
| 30.6338 | 30.3838 | 0.99221 |  | 27.5 | 27.25 | 0.99029 | 29.7615 | 29.5115 | 0.99365 |  | 0.94673 | 0.94355
| 31.1338 | 30.8838 | 0.98441 |  | 28 | 27.75 | 0.99728 | 30.2615 | 30.0115 | 0.9783 |  | 0.95341 | 0.92898
| 31.6338 | 31.3838 | 0.97806 |  | 28.5 | 28.25 | 0.99853 | 30.7615 | 30.5115 | 0.97514 |  | 0.9546 | 0.92598
| 32.1338 | 31.8838 | 0.96836 |  | 29 | 28.75 | 1 | 31.2615 | 31.0115 | 0.96957 |  | 0.95601 | 0.92069
| 32.6338 | 32.3838 | 0.9568 |  | 29.5 | 29.25 | 0.99762 | 31.7615 | 31.5115 | 0.96996 |  | 0.95374 | 0.92106
| 33.1338 | 32.8838 | 0.94512 |  | 30 | 29.75 | 0.99495 | 32.2615 | 32.0115 | 0.95585 |  | 0.95118 | 0.90766
| 33.6338 | 33.3838 | 0.93803 |  | 30.5 | 30.25 | 0.9895 | 32.7615 | 32.5115 | 0.95227 |  | 0.94597 | 0.90426
| 34.1338 | 33.8838 | 0.92562 |  | 31 | 30.75 | 0.98482 | 33.2615 | 33.0115 | 0.95032 |  | 0.9415 | 0.90241
| 34.6338 | 34.3838 | 0.91167 |  | 31.5 | 31.25 | 0.97743 | 33.7615 | 33.5115 | 0.9336 |  | 0.93444 | 0.88653
| 35.1338 | 34.8838 | 0.90704 |  | 32 | 31.75 | 0.97074 | 34.2615 | 34.0115 | 0.9269 |  | 0.92804 | 0.88017
| 35.6338 | 35.3838 | 0.88747 |  | 32.5 | 32.25 | 0.96113 | 34.7615 | 34.5115 | 0.9027 |  | 0.91885 | 0.85719
| 36.1338 | 35.8838 | 0.88227 |  | 33 | 32.75 | 0.95196 | 35.2615 | 35.0115 | 0.89767 |  | 0.91009 | 0.85242
| 36.6338 | 36.3838 | 0.86155 |  | 33.5 | 33.25 | 0.94105 | 35.7615 | 35.5115 | 0.88809 |  | 0.89965 | 0.84332
| 37.1338 | 36.8838 | 0.84294 |  | 34 | 33.75 | 0.92994 | 36.2615 | 36.0115 | 0.86058 |  | 0.88903 | 0.81719
| 37.6338 | 37.3838 | 0.83073 |  | 34.5 | 34.25 | 0.91732 | 36.7615 | 36.5115 | 0.87106 |  | 0.87696 | 0.82715
| 38.1338 | 37.8838 | 0.8238 |  | 35 | 34.75 | 0.9057 | 37.2615 | 37.0115 | 0.84335 |  | 0.86585 | 0.80084
| 38.6338 | 38.3838 | 0.80319 |  | 35.5 | 35.25 | 0.89207 | 37.7615 | 37.5115 | 0.83731 |  | 0.85282 | 0.7951
| 39.1338 | 38.8838 | 0.78099 |  | 36 | 35.75 | 0.87869 | 38.2615 | 38.0115 | 0.82929 |  | 0.84004 | 0.78748
| 39.6338 | 39.3838 | 0.76967 |  | 36.5 | 36.25 | 0.8653 | 38.7615 | 38.5115 | 0.80248 |  | 0.82723 | 0.76202
| 40.1338 | 39.8838 | 0.75601 |  | 37 | 36.75 | 0.85175 | 39.2615 | 39.0115 | 0.78401 |  | 0.81428 | 0.74448
| 40.6338 | 40.3838 | 0.74063 |  | 37.5 | 37.25 | 0.83719 | 39.7615 | 39.5115 | 0.77033 |  | 0.80036 | 0.7315
| 41.1338 | 40.8838 | 0.72372 |  | 38 | 37.75 | 0.82325 | 40.2615 | 40.0115 | 0.7586 |  | 0.78704 | 0.72036
| 41.6338 | 41.3838 | 0.70713 |  | 38.5 | 38.25 | 0.80829 | 40.7615 | 40.5115 | 0.75225 |  | 0.77273 | 0.71433
| 42.1338 | 41.8838 | 0.68877 |  | 39 | 38.75 | 0.79396 | 41.2615 | 41.0115 | 0.73187 |  | 0.75904 | 0.69497
| 42.6338 | 42.3838 | 0.67726 |  | 39.5 | 39.25 | 0.77896 | 41.7615 | 41.5115 | 0.71578 |  | 0.74469 | 0.67969
| 43.1338 | 42.8838 | 0.66011 |  | 40 | 39.75 | 0.76442 | 42.2615 | 42.0115 | 0.70592 |  | 0.73079 | 0.67033
| 43.6338 | 43.3838 | 0.6421 |  | 40.5 | 40.25 | 0.74962 | 42.7615 | 42.5115 | 0.68636 |  | 0.71665 | 0.65176
| 44.1338 | 43.8838 | 0.63376 |  | 41 | 40.75 | 0.73476 | 43.2615 | 43.0115 | 0.66037 |  | 0.70244 | 0.62708
| 44.6338 | 44.3838 | 0.61224 |  | 41.5 | 41.25 | 0.71956 | 43.7615 | 43.5115 | 0.65624 |  | 0.68791 | 0.62315
| 45.1338 | 44.8838 | 0.59722 |  | 42 | 41.75 | 0.70517 | 44.2615 | 44.0115 | 0.64073 |  | 0.67415 | 0.60843
| 45.6338 | 45.3838 | 0.58635 |  | 42.5 | 42.25 | 0.69026 | 44.7615 | 44.5115 | 0.62371 |  | 0.65989 | 0.59227
| 46.1338 | 45.8838 | 0.57181 |  | 43 | 42.75 | 0.67561 | 45.2615 | 45.0115 | 0.61906 |  | 0.64589 | 0.58785
| 46.6338 | 46.3838 | 0.55121 |  | 43.5 | 43.25 | 0.6612 | 45.7615 | 45.5115 | 0.60305 |  | 0.63211 | 0.57265
| 47.1338 | 46.8838 | 0.53741 |  | 44 | 43.75 | 0.64675 | 46.2615 | 46.0115 | 0.58278 |  | 0.6183 | 0.5534
| 47.6338 | 47.3838 | 0.52364 |  | 44.5 | 44.25 | 0.63252 | 46.7615 | 46.5115 | 0.57939 |  | 0.60469 | 0.55018
| 48.1338 | 47.8838 | 0.50408 |  | 45 | 44.75 | 0.61808 | 47.2615 | 47.0115 | 0.55929 |  | 0.59089 | 0.53109
| 48.6338 | 48.3838 | 0.49254 |  | 45.5 | 45.25 | 0.60384 | 47.7615 | 47.5115 | 0.54421 |  | 0.57727 | 0.51677
| 49.1338 | 48.8838 | 0.48139 |  | 46 | 45.75 | 0.59007 | 48.2615 | 48.0115 | 0.53489 |  | 0.56412 | 0.50793
| 49.6338 | 49.3838 | 0.45796 |  | 46.5 | 46.25 | 0.57579 | 48.7615 | 48.5115 | 0.5165 |  | 0.55046 | 0.49046
| 50.1338 | 49.8838 | 0.4514 |  | 47 | 46.75 | 0.56172 | 49.2615 | 49.0115 | 0.50255 |  | 0.53701 | 0.47722
| 50.6338 | 50.3838 | 0.43487 |  | 47.5 | 47.25 | 0.54825 | 49.7615 | 49.5115 | 0.4971 |  | 0.52413 | 0.47204
| 51.1338 | 50.8838 | 0.42462 |  | 48 | 47.75 | 0.53456 | 50.2615 | 50.0115 | 0.4736 |  | 0.51105 | 0.44972
| 51.6338 | 51.3838 | 0.40736 |  | 48.5 | 48.25 | 0.52107 | 50.7615 | 50.5115 | 0.46639 |  | 0.49815 | 0.44288
| 52.1338 | 51.8838 | 0.38887 |  | 49 | 48.75 | 0.50813 | 51.2615 | 51.0115 | 0.4519 |  | 0.48578 | 0.42911
| 52.6338 | 52.3838 | 0.38406 |  | 49.5 | 49.25 | 0.49484 | 51.7615 | 51.5115 | 0.43608 |  | 0.47308 | 0.41409
| 53.1338 | 52.8838 | 0.36598 |  | 50 | 49.75 | 0.48178 | 52.2615 | 52.0115 | 0.42602 |  | 0.46059 | 0.40454
| 53.6338 | 53.3838 | 0.35747 |  | 50.5 | 50.25 | 0.46886 | 52.7615 | 52.5115 | 0.41905 |  | 0.44824 | 0.39792
| 54.1338 | 53.8838 | 0.33936 |  | 51 | 50.75 | 0.45597 | 53.2615 | 53.0115 | 0.39863 |  | 0.43591 | 0.37853
| 54.6338 | 54.3838 | 0.32775 |  | 51.5 | 51.25 | 0.4435 | 53.7615 | 53.5115 | 0.38791 |  | 0.42399 | 0.36836
| 55.1338 | 54.8838 | 0.31584 |  | 52 | 51.75 | 0.43081 | 54.2615 | 54.0115 | 0.38222 |  | 0.41186 | 0.36295
| 55.6338 | 55.3838 | 0.30218 |  | 52.5 | 52.25 | 0.41845 | 54.7615 | 54.5115 | 0.35709 |  | 0.40004 | 0.33909
| 56.1338 | 55.8838 | 0.28624 |  | 53 | 52.75 | 0.40642 | 55.2615 | 55.0115 | 0.35109 |  | 0.38854 | 0.33339
| 56.6338 | 56.3838 | 0.27651 |  | 53.5 | 53.25 | 0.39408 | 55.7615 | 55.5115 | 0.33297 |  | 0.37675 | 0.31618
| 57.1338 | 56.8838 | 0.2658 |  | 54 | 53.75 | 0.38191 | 56.2615 | 56.0115 | 0.32412 |  | 0.36511 | 0.30778
| 57.6338 | 57.3838 | 0.25382 |  | 54.5 | 54.25 | 0.37027 | 56.7615 | 56.5115 | 0.30994 |  | 0.35398 | 0.29432
| 58.1338 | 57.8838 | 0.24171 |  | 55 | 54.75 | 0.35842 | 57.2615 | 57.0115 | 0.30951 |  | 0.34266 | 0.29391
| 58.6338 | 58.3838 | 0.2303 |  | 55.5 | 55.25 | 0.34657 | 57.7615 | 57.5115 | 0.29318 |  | 0.33132 | 0.2784
| 59.1338 | 58.8838 | 0.21488 |  | 56 | 55.75 | 0.33507 | 58.2615 | 58.0115 | 0.27916 |  | 0.32033 | 0.26508
| 59.6338 | 59.3838 | 0.20197 |  | 56.5 | 56.25 | 0.32358 | 58.7615 | 58.5115 | 0.2691 |  | 0.30934 | 0.25554
| 60.1338 | 59.8838 | 0.19254 |  | 57 | 56.75 | 0.31253 | 59.2615 | 59.0115 | 0.25324 |  | 0.29878 | 0.24048
| 60.6338 | 60.3838 | 0.18202 |  | 57.5 | 57.25 | 0.30105 | 59.7615 | 59.5115 | 0.24748 |  | 0.2878 | 0.235
| 61.1338 | 60.8838 | 0.16906 |  | 58 | 57.75 | 0.28987 | 60.2615 | 60.0115 | 0.23688 |  | 0.27712 | 0.22494
| 61.6338 | 61.3838 | 0.15923 |  | 58.5 | 58.25 | 0.27887 | 60.7615 | 60.5115 | 0.22336 |  | 0.2666 | 0.2121
| 62.1338 | 61.8838 | 0.14845 |  | 59 | 58.75 | 0.26759 | 61.2615 | 61.0115 | 0.21389 |  | 0.25582 | 0.2031
| 62.6338 | 62.3838 | 0.13911 |  | 59.5 | 59.25 | 0.25644 | 61.7615 | 61.5115 | 0.20013 |  | 0.24516 | 0.19004
| 63.1338 | 62.8838 | 0.12716 |  | 60 | 59.75 | 0.2458 | 62.2615 | 62.0115 | 0.18992 |  | 0.23498 | 0.18035
| 63.6338 | 63.3838 | 0.11375 |  | 60.5 | 60.25 | 0.23467 | 62.7615 | 62.5115 | 0.17348 |  | 0.22435 | 0.16473
| 64.1338 | 63.8838 | 0.10372 |  | 61 | 60.75 | 0.22368 | 63.2615 | 63.0115 | 0.16752 |  | 0.21384 | 0.15907
| 64.6338 | 64.3838 | 0.093689 |  | 61.5 | 61.25 | 0.21302 | 63.7615 | 63.5115 | 0.15594 |  | 0.20365 | 0.14808
| 65.1338 | 64.8838 | 0.083824 |  | 62 | 61.75 | 0.20211 | 64.2615 | 64.0115 | 0.13907 |  | 0.19322 | 0.13206
| 65.6338 | 65.3838 | 0.072444 |  | 62.5 | 62.25 | 0.19126 | 64.7615 | 64.5115 | 0.1298 |  | 0.18285 | 0.12325
| 66.1338 | 65.8838 | 0.065213 |  | 63 | 62.75 | 0.18042 | 65.2615 | 65.0115 | 0.1176 |  | 0.17248 | 0.11167
| 66.6338 | 66.3838 | 0.052587 |  | 63.5 | 63.25 | 0.16956 | 65.7615 | 65.5115 | 0.10264 |  | 0.1621 | 0.097464
| 67.1338 | 66.8838 | 0.045126 |  | 64 | 63.75 | 0.15902 | 66.2615 | 66.0115 | 0.092364 |  | 0.15202 | 0.087707
| 67.6338 | 67.3838 | 0.035415 |  | 64.5 | 64.25 | 0.14771 | 66.7615 | 66.5115 | 0.083661 |  | 0.14121 | 0.079444
| 68.1338 | 67.8838 | 0.027131 |  | 65 | 64.75 | 0.1367 | 67.2615 | 67.0115 | 0.071231 |  | 0.13069 | 0.06764
| 68.8067 | 68.3838 | 0.017907 |  | 65.5 | 65.25 | 0.12562 | 67.7615 | 67.5115 | 0.059034 |  | 0.12009 | 0.056058
| 69.4796 | 69.2296 | 0.0062697 |  | 66 | 65.75 | 0.11348 | 68.2615 | 68.0115 | 0.043175 |  | 0.10849 | 0.040998
[5 empty] | 66.5 | 66.25 | 0.1014 | 68.7615 | 68.5115 | 0.02716 |  | 0.096935 | 0.02579
[5 empty] | 67 | 66.75 | 0.089931 | 69.26 | 69.0115 | 0.012236 |  | 0.085975 | 0.011619
[5 empty] | 67.5 | 67.25 | 0.078238 |  |  |  |  | 0.074796 | 0
[5 empty] | 68 | 67.75 | 0.066477 |  |  |  |  | 0.063553 | 0
[5 empty] | 68.5 | 68.25 | 0.046065 |  |  |  |  | 0.044038 | 0
[5 empty] | 69 | 68.75 | 0.022724 |  |  |  |  | 0.021724 | 0
[5 empty] | 69.5 | 69.25 | 0 |  |  |  |  | 0 | 0