## Description
General information
Catalogue version: 6 October 2021
Created by Jaroslav Šolc, jsolc@cmi.cz, Czech Metrology Institute
This Excel file contains the catalogue of X-ray energy spectra of electronic brachytherapy (eBT) device Papillon50 manufactured by Ariane Medical Systems.
The catalogue was compiled from suitable data available in literature and from the data measured or created by the project "18NRM02 PRISM-eBT" (Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices; http://www.ebt-empir.eu/) of the European Metrology Programme for Innovation and Research (EMPIR).
The main goal of the catalogue was to develop a set of harmonised reference X-ray qualities denoted as “eBT equivalent” that would be established for in-laboratory activities by the X-ray tubes usually available for calibration in calibration laboratories, but the tube voltage and the filter are specifically selected to simulate the energy spectrum of the corresponding eBT radiation as closely as possible at the reference point.
In general, the catalogue consists of:
1) X-ray photon fluence eBT spectra measured and/or Monte Carlo simulated in air and/or in water, whenever they were available;
2) eBT-equivalent X-ray photon fluence spectra that match selected eBT spectra. Typically, the eBT-equivalent spectra were calculated by a dedicated software, usually SpekPy software. If available, the eBT-equivalent spectra already realized are also included. Please note that the eBT-equivalent spectra may be site-dependent, based on properties of X-ray generator available (anode angle, inherent filtration) and the distance from anode to measurement point (attenuation in air). Realized eBT-equivalent spectra should be validated by spectrometry.
Specifically, this catalog includes spectra listed below:
Sheet | Description
MC, water, 1 cm, eBT-eq. | 1) MC simulated spectra in 1 cm water depth obtained from the measured spectra in air presented in the other two sheets 2) eBT-equivalent spectra at 1 cm water depth
EXP, air, 25 cm | Measured spectra in air, 25 cm source-to-detector distance
EXP, air, 532 cm | Measured spectrum in air, 532 cm source-to-detector distance

## MC, water, 1 cm, eBT-eq.
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
3 June 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Papillon50 |  |  | This sheet also contains eBT-equivalent spectra
Manufacturer: | Ariane Medical Systems
Spectra information
origin | MC simulation - source spectra taken from the measurement at 25 cm and at 532 cm and recalculated to 0 cm (see sheets "EXP, air, 25 cm" and "EXP, air, 532 cm")
Monte Carlo (MC) code | MCNP6.2
MC simulation settings | EPRDATA14 photon library
authors | CEA (measurement at 532 cm), CMI + AU (masurements at 25 cm, MC simulations)
location | -
publication | unpublished
DOI | N/A
year of the result | 2021
material | water
source-to-detector distance (cm) | 1 | depth in water; a water phantom (cube) is pushed tightly towards the applicator
description of source-to-detector distance | on device axis, distance from water surface - a water phantom (cube) is pushed tightly towards the applicator
applicator | diameter 22 mm
other information | Papillon50 device not simulated. The source spectrum was taken from the measurements at 25 cm and 532 cm with 22 mm applicator (and recalculated to 0 cm, see sheets "EXP, air, 25 cm" and "EXP, air, 532 cm")
Device setting
nominal bias voltage (kV) | 50 | nominal, but the real voltage was obviously lower in CMI/AU measurements, see plot
bias voltage sigma (kV) | N/A
current (mA) | N/A |  |  |  |  | A simple visualization of Monte Carlo model: point cone beam (45° angle) impiging on water surface. Point of interest is located on beam axis 1 cm from water surface
input parameters for calculation of eBT-equivalent spectra
Method | [7 empty] | SpekPy fit | SpekPy fit | SpekPy fit | [5 empty] | SpekPy fit
Anode material | [7 empty] | W | W | W | [5 empty] | W
Anode angle (degree) | [7 empty] | 21 | 21 | 21 | [5 empty] | 21
X-ray generator voltage (kWp) | [7 empty] | 47.8 | 47.8 | 47.8 | [5 empty] | 50
Filtration 1 - material and thickness (mm) | Be | [6 empty] | 1 | 1 | 1 | [5 empty] | 1
Filtration 2 - material and thickness (mm) | Air | [6 empty] | 964 | 964 | 964 | [5 empty] | 964
Filtration 3 - material and thickness (mm) | Al | [6 empty] | 1.472 | 1.476 | 1.484 | [5 empty] | 1.134
characteristics of eBT and eBT-equivalent spectra
Filtration material, maximum thickness (mm; user-defined) and density (g/cm3) | Al | 2 | 2.7 | 2.7 | 2.7 |  |  | 2.7 | 2.7 | 2.7 |  |  | 2.7 |  |  | 2.7
Emean (in F; keV) |  |  | 29.2802 | 29.2773 | 29.2744 |  |  | 29.3715 | 29.3606 | 29.389 |  |  | 29.3117 |  |  | 29.2263 | average photon energy
Emean (in Kair; keV) |  |  | 25.4963 | 25.5071 | 25.5249 |  |  | 25.5181 | 25.5031 | 25.542 |  |  | 24.7371 |  |  | 24.7105 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 23.8131 | 23.8291 | 23.8545 |  |  | 23.8088 | 23.793 | 23.8339 |  |  | 22.8129 |  |  | 22.8087 | energy of monoenergetic photons with the same HVL as the given spectrum | 0.01217 | 0.010239 | 0.010239 | 0.010239
1st HVL in Al - Excel (in Kair; mm) |  |  | 1.2263 | 1.2286 | 1.2323 |  |  | 1.2257 | 1.2234 | 1.2293 |  |  | 1.0875 |  |  | 1.0869 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value; determined by Excel script | 53982835.6143 | 54286198.9423 | 0.061974 | 72459856.0785
1st HVL in Al - Method (in Kair; mm) |  |  | 1.219 | 1.222 | 1.225 |  |  | 1.219 | 1.222 | 1.225 |  |  | 1.083 |  |  | 1.083 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value, determined by SpekPy fit
| based on measurements in air in 25 cm (see sheet "EXP, air, 25 cm") | [9 empty] | based on measurements in air in 532 cm (see sheet "EXP, air, 532 cm")
| eBT spectra in 1 cm water depth |  |  |  |  | eBT-equivalent spectra |  |  |  |  | eBT spectrum in 1 cm water depth |  |  |  | eBT-equivalent spectrum
applicator diameter (mm) |  |  | 22 | 25 | 30 |  |  | 22 | 25 | 30 |  |  | 22 |  |  | 22
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Fluence (relative) | Fluence (relative) | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Fluence (relative) | Fluence (relative) | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative)
| 3 |  |  |  |  | 1.1 |  |  |  |  | 3.1 |  |  | 1.1
| 3.1 | 3.05 | 1.0511e-09 | 8.7512e-10 | 6.5594e-10 | 1.2 | 1.15 | 0 | 0 | 0 | 3.3 | 3.2 | 1.8713e-08 | 1.2 | 1.15 | 0 |  | 0 | 0 | 3.0916e-09 | 0
| 3.2 | 3.15 | 6.9596e-10 | 6.2607e-10 | 2.1421e-10 | 1.3 | 1.25 | 0 | 0 | 0 | 3.5 | 3.4 | 9.1039e-09 | 1.3 | 1.25 | 0 |  | 0 | 0 | 1.5041e-09 | 0
| 3.3 | 3.25 | 3.7515e-10 | 1.6088e-09 | 1.2375e-09 | 1.4 | 1.35 | 0 | 0 | 0 | 3.7 | 3.6 | 2.3765e-08 | 1.4 | 1.35 | 0 |  | 0 | 0 | 3.9263e-09 | 0
| 3.4 | 3.35 | 1.1236e-09 | 7.7292e-10 | 6.0089e-10 | 1.5 | 1.45 | 0 | 0 | 0 | 3.9 | 3.8 | 1.0751e-08 | 1.5 | 1.45 | 0 |  | 0 | 0 | 1.7763e-09 | 0
| 3.5 | 3.45 | 2.8952e-09 | 2.3642e-09 | 1.7552e-09 | 1.6 | 1.55 | 0 | 0 | 0 | 4.1 | 4 | 1.398e-08 | 1.6 | 1.55 | 0 |  | 0 | 0 | 2.3097e-09 | 0
| 3.6 | 3.55 | 7.6478e-12 | 6.4148e-12 | 6.5505e-11 | 1.7 | 1.65 | 0 | 0 | 0 | 4.3 | 4.2 | 1.3562e-08 | 1.7 | 1.65 | 0 |  | 0 | 0 | 2.2407e-09 | 0
| 3.7 | 3.65 | 2.5336e-10 | 2.1095e-10 | 2.3214e-10 | 1.8 | 1.75 | 0 | 0 | 0 | 4.5 | 4.4 | 1.2226e-08 | 1.8 | 1.75 | 0 |  | 0 | 0 | 2.0199e-09 | 0
| 3.8 | 3.75 | 5.3307e-11 | 2.7787e-10 | 1.4811e-09 | 1.9 | 1.85 | 0 | 0 | 0 | 4.7 | 4.6 | 1.5952e-08 | 1.9 | 1.85 | 0 |  | 0 | 0 | 2.6355e-09 | 0
| 3.9 | 3.85 | 2.2809e-10 | 0 | 0 | 2 | 1.95 | 0 | 0 | 0 | 4.9 | 4.8 | 1.8148e-08 | 2 | 1.95 | 0 |  | 0 | 0 | 2.9983e-09 | 0
| 4 | 3.95 | 6.6914e-09 | 5.8997e-09 | 4.4162e-09 | 2.1 | 2.05 | 0 | 0 | 0 | 5.1 | 5 | 3.5832e-08 | 2.1 | 2.05 | 0 |  | 0 | 0 | 5.9199e-09 | 0
| 4.1 | 4.05 | 1.251e-09 | 1.042e-09 | 7.8148e-10 | 2.2 | 2.15 | 0 | 0 | 0 | 5.3 | 5.2 | 5.7235e-09 | 2.2 | 2.15 | 0 |  | 0 | 0 | 9.456e-10 | 0
| 4.2 | 4.15 | 5.9566e-09 | 4.9638e-09 | 3.4387e-09 | 2.3 | 2.25 | 0 | 0 | 0 | 5.5 | 5.4 | 1.2287e-08 | 2.3 | 2.25 | 0 |  | 0 | 0 | 2.0299e-09 | 0
| 4.3 | 4.25 | 4.8034e-11 | 3.9995e-11 | 4.897e-10 | 2.4 | 2.35 | 0 | 0 | 0 | 5.7 | 5.6 | 2.0426e-08 | 2.4 | 2.35 | 0 |  | 0 | 0 | 3.3747e-09 | 0
| 4.4 | 4.35 | 4.8816e-09 | 4.2024e-09 | 2.6553e-09 | 2.5 | 2.45 | 0 | 0 | 0 | 5.9 | 5.8 | 9.6924e-09 | 2.5 | 2.45 | 0 |  | 0 | 0 | 1.6013e-09 | 0
| 4.5 | 4.45 | 3.5747e-09 | 2.9694e-09 | 2.2466e-09 | 2.6 | 2.55 | 0 | 0 | 0 | 6.1 | 6 | 3.9574e-08 | 2.6 | 2.55 | 0 |  | 0 | 0 | 6.5381e-09 | 0
| 4.6 | 4.55 | 1.3939e-09 | 1.6076e-09 | 8.6502e-10 | 2.7 | 2.65 | 0 | 0 | 0 | 6.3 | 6.2 | 2.3311e-08 | 2.7 | 2.65 | 0 |  | 0 | 0 | 3.8512e-09 | 0
| 4.7 | 4.65 | 2.3578e-09 | 4.677e-09 | 3.5373e-09 | 2.8 | 2.75 | 0 | 0 | 0 | 6.5 | 6.4 | 8.4107e-09 | 2.8 | 2.75 | 0 |  | 0 | 0 | 1.3896e-09 | 0
| 4.8 | 4.75 | 2.4807e-09 | 2.0654e-09 | 2.6318e-09 | 2.9 | 2.85 | 0 | 0 | 0 | 6.7 | 6.6 | 2.9971e-08 | 2.9 | 2.85 | 0 |  | 0 | 0 | 4.9516e-09 | 0
| 4.9 | 4.85 | 2.024e-09 | 1.6954e-09 | 1.1849e-09 | 3 | 2.95 | 0 | 0 | 0 | 6.9 | 6.8 | 2.978e-08 | 3 | 2.95 | 0 |  | 0 | 0 | 4.9201e-09 | 0
| 5 | 4.95 | 2.2233e-09 | 1.9439e-09 | 3.0488e-09 | 3.1 | 3.05 | 0 | 0 | 0 | 7.1 | 7 | 2.6722e-08 | 3.1 | 3.05 | 0 |  | 0 | 0 | 4.4149e-09 | 0
| 5.1 | 5.05 | 1.4852e-09 | 1.2578e-09 | 1.2936e-09 | 3.2 | 3.15 | 0 | 0 | 0 | 7.3 | 7.2 | 4.4625e-08 | 3.2 | 3.15 | 0 |  | 0 | 0 | 7.3726e-09 | 0
| 5.2 | 5.15 | 2.6727e-09 | 3.4927e-09 | 2.5886e-09 | 3.3 | 3.25 | 0 | 0 | 0 | 7.5 | 7.4 | 4.2357e-08 | 3.3 | 3.25 | 0 |  | 0 | 0 | 6.998e-09 | 0
| 5.3 | 5.25 | 1.8586e-09 | 2.7743e-09 | 2.639e-09 | 3.4 | 3.35 | 0 | 0 | 0 | 7.7 | 7.6 | 3.5908e-08 | 3.4 | 3.35 | 0 |  | 0 | 0 | 5.9325e-09 | 0
| 5.4 | 5.35 | 7.1e-09 | 4.6565e-09 | 3.4899e-09 | 3.5 | 3.45 | 0 | 0 | 0 | 7.9 | 7.8 | 4.9664e-08 | 3.5 | 3.45 | 0 |  | 0 | 0 | 8.2052e-09 | 0
| 5.5 | 5.45 | 9.4674e-09 | 5.7787e-09 | 4.0969e-09 | 3.6 | 3.55 | 0 | 0 | 0 | 8.1 | 8 | 4.6042e-08 | 3.6 | 3.55 | 0 |  | 0 | 0 | 7.6068e-09 | 0
| 5.6 | 5.55 | 4.4873e-09 | 3.6363e-09 | 2.727e-09 | 3.7 | 3.65 | 0 | 0 | 0 | 8.3 | 8.2 | 5.621e-08 | 3.7 | 3.65 | 0 |  | 0 | 0 | 9.2866e-09 | 0
| 5.7 | 5.65 | 5.3392e-09 | 3.9555e-09 | 2.801e-09 | 3.8 | 3.75 | 0 | 0 | 0 | 8.5 | 8.4 | 8.3274e-08 | 3.8 | 3.75 | 0 |  | 0 | 0 | 1.3758e-08 | 0
| 5.8 | 5.75 | 1.9107e-09 | 2.1651e-09 | 2.6898e-09 | 3.9 | 3.85 | 0 | 0 | 0 | 8.7 | 8.6 | 2.3048e-07 | 3.9 | 3.85 | 0 |  | 0 | 0 | 3.8078e-08 | 0
| 5.9 | 5.85 | 2.3626e-09 | 2.909e-09 | 2.2759e-09 | 4 | 3.95 | 0 | 0 | 0 | 8.9 | 8.8 | 4.1598e-07 | 4 | 3.95 | 0 |  | 0 | 0 | 6.8725e-08 | 0
| 6 | 5.95 | 5.0192e-10 | 3.1283e-10 | 1.301e-10 | 4.1 | 4.05 | 0 | 0 | 0 | 9.1 | 9 | 7.3453e-07 | 4.1 | 4.05 | 0 |  | 0 | 0 | 1.2135e-07 | 0
| 6.1 | 6.05 | 8.9858e-09 | 7.5697e-09 | 5.8029e-09 | 4.2 | 4.15 | 0 | 0 | 0 | 9.3 | 9.2 | 1.4729e-06 | 4.2 | 4.15 | 0 |  | 0 | 0 | 2.4334e-07 | 0
| 6.2 | 6.15 | 2.7831e-09 | 2.3156e-09 | 1.7329e-09 | 4.3 | 4.25 | 0 | 0 | 0 | 9.5 | 9.4 | 3.3562e-06 | 4.3 | 4.25 | 0 |  | 0 | 0 | 5.5449e-07 | 0
| 6.3 | 6.25 | 5.1591e-09 | 1.1571e-08 | 8.8493e-09 | 4.4 | 4.35 | 0 | 0 | 0 | 9.7 | 9.6 | 4.183e-06 | 4.4 | 4.35 | 0 |  | 0 | 0 | 6.9109e-07 | 0
| 6.4 | 6.35 | 8.433e-09 | 4.252e-09 | 5.3301e-09 | 4.5 | 4.45 | 0 | 0 | 0 | 9.9 | 9.8 | 7.7848e-06 | 4.5 | 4.45 | 0 |  | 0 | 0 | 1.2862e-06 | 0
| 6.5 | 6.45 | 1.4271e-08 | 1.1897e-08 | 8.7764e-09 | 4.6 | 4.55 | 0 | 0 | 0 | 10.1 | 10 | 1.2952e-05 | 4.6 | 4.55 | 0 |  | 0 | 0 | 2.1398e-06 | 0
| 6.6 | 6.55 | 1.3849e-09 | 1.2119e-09 | 7.9663e-10 | 4.7 | 4.65 | 0 | 0 | 0 | 10.3 | 10.2 | 3.1685e-05 | 4.7 | 4.65 | 0 |  | 0 | 0 | 5.2348e-06 | 0
| 6.7 | 6.65 | 5.269e-09 | 6.242e-09 | 4.7363e-09 | 4.8 | 4.75 | 0 | 0 | 0 | 10.5 | 10.4 | 4.2163e-05 | 4.8 | 4.75 | 0 |  | 0 | 0 | 6.9659e-06 | 0
| 6.8 | 6.75 | 2.5606e-09 | 2.1362e-09 | 1.6136e-09 | 4.9 | 4.85 | 0 | 0 | 0 | 10.7 | 10.6 | 6.0183e-05 | 4.9 | 4.85 | 0 |  | 0 | 0 | 9.9429e-06 | 0
| 6.9 | 6.85 | 4.1287e-09 | 3.6827e-09 | 2.5561e-09 | 5 | 4.95 | 0 | 0 | 0 | 10.9 | 10.8 | 5.4797e-05 | 5 | 4.95 | 0 |  | 0 | 0 | 9.0532e-06 | 0
| 7 | 6.95 | 1.3231e-09 | 2.8255e-09 | 9.8698e-10 | 5.1 | 5.05 | 0 | 0 | 0 | 11.1 | 11 | 2.4073e-05 | 5.1 | 5.05 | 0 |  | 0 | 0 | 3.9772e-06 | 0
| 7.1 | 7.05 | 7.0228e-09 | 5.9038e-09 | 3.7824e-09 | 5.2 | 5.15 | 0 | 0 | 0 | 11.3 | 11.2 | 3.2673e-05 | 5.2 | 5.15 | 0 |  | 0 | 0 | 5.3981e-06 | 0
| 7.2 | 7.15 | 1.0122e-08 | 7.988e-09 | 6.1607e-09 | 5.3 | 5.25 | 0 | 0 | 0 | 11.5 | 11.4 | 4.4424e-05 | 5.3 | 5.25 | 0 |  | 0 | 0 | 7.3394e-06 | 0
| 7.3 | 7.25 | 7.6846e-09 | 6.1556e-09 | 5.1459e-09 | 5.4 | 5.35 | 0 | 0 | 0 | 11.7 | 11.6 | 6.7015e-05 | 5.4 | 5.35 | 0 |  | 0 | 0 | 1.1072e-05 | 0
| 7.4 | 7.35 | 1.2043e-08 | 1.1234e-08 | 8.5476e-09 | 5.5 | 5.45 | 0 | 0 | 0 | 11.9 | 11.8 | 8.2421e-05 | 5.5 | 5.45 | 0 |  | 0 | 0 | 1.3617e-05 | 0
| 7.5 | 7.45 | 7.3432e-09 | 6.1139e-09 | 4.5236e-09 | 5.6 | 5.55 | 0 | 0 | 0 | 12.1 | 12 | 0.00010814 | 5.6 | 5.55 | 0 |  | 0 | 0 | 1.7866e-05 | 0
| 7.6 | 7.55 | 9.3292e-09 | 9.6258e-09 | 7.9206e-09 | 5.7 | 5.65 | 0 | 0 | 0 | 12.3 | 12.2 | 0.00014481 | 5.7 | 5.65 | 0 |  | 0 | 0 | 2.3925e-05 | 0
| 7.7 | 7.65 | 8.8917e-09 | 6.7154e-09 | 4.9504e-09 | 5.8 | 5.75 | 0 | 0 | 0 | 12.5 | 12.4 | 0.0001321 | 5.8 | 5.75 | 0 |  | 0 | 0 | 2.1825e-05 | 0
| 7.8 | 7.75 | 4.2792e-09 | 3.2444e-09 | 5.5895e-09 | 5.9 | 5.85 | 0 | 0 | 0 | 12.7 | 12.6 | 0.0001211 | 5.9 | 5.85 | 0 |  | 0 | 0 | 2.0007e-05 | 0
| 7.9 | 7.85 | 8.1272e-09 | 6.8765e-09 | 5.0539e-09 | 6 | 5.95 | 0 | 0 | 0 | 12.9 | 12.8 | 0.00013802 | 6 | 5.95 | 0 |  | 0 | 0 | 2.2802e-05 | 0
| 8 | 7.95 | 3.4362e-09 | 3.299e-09 | 2.0517e-09 | 6.1 | 6.05 | 0 | 0 | 0 | 13.1 | 13 | 0.000135 | 6.1 | 6.05 | 0 |  | 0 | 0 | 2.2304e-05 | 0
| 8.1 | 8.05 | 6.4011e-09 | 5.3505e-09 | 1.7146e-09 | 6.2 | 6.15 | 0 | 0 | 0 | 13.3 | 13.2 | 0.0001741 | 6.2 | 6.15 | 0 |  | 0 | 0 | 2.8763e-05 | 0
| 8.2 | 8.15 | 3.5505e-09 | 3.0455e-09 | 2.3119e-09 | 6.3 | 6.25 | 0 | 0 | 0 | 13.5 | 13.4 | 0.00021502 | 6.3 | 6.25 | 0 |  | 0 | 0 | 3.5524e-05 | 0
| 8.3 | 8.25 | 1.0204e-09 | 8.3429e-10 | 5.7844e-10 | 6.4 | 6.35 | 0 | 0 | 0 | 13.7 | 13.6 | 0.00026202 | 6.4 | 6.35 | 0 |  | 0 | 0 | 4.3289e-05 | 0
| 8.4 | 8.35 | 7.5806e-09 | 6.5605e-09 | 4.1935e-09 | 6.5 | 6.45 | 0 | 0 | 0 | 13.9 | 13.8 | 0.00031481 | 6.5 | 6.45 | 0 |  | 0 | 0 | 5.201e-05 | 0
| 8.5 | 8.45 | 5.4601e-09 | 5.2865e-09 | 3.7834e-09 | 6.6 | 6.55 | 0 | 0 | 0 | 14.1 | 14 | 0.00036548 | 6.6 | 6.55 | 0 |  | 0 | 0 | 6.0383e-05 | 0
| 8.6 | 8.55 | 2.6709e-09 | 1.4574e-09 | 1.0604e-09 | 6.7 | 6.65 | 0 | 0 | 0 | 14.3 | 14.2 | 0.00040041 | 6.7 | 6.65 | 0 |  | 0 | 0 | 6.6153e-05 | 0
| 8.7 | 8.65 | 1.634e-09 | 1.6507e-09 | 8.7684e-10 | 6.8 | 6.75 | 0 | 0 | 0 | 14.5 | 14.4 | 0.00046159 | 6.8 | 6.75 | 0 |  | 0 | 0 | 7.6261e-05 | 0
| 8.8 | 8.75 | 1.2973e-08 | 6.6958e-09 | 4.7792e-09 | 6.9 | 6.85 | 0 | 0 | 0 | 14.7 | 14.6 | 0.00051223 | 6.9 | 6.85 | 0 |  | 0 | 0 | 8.4627e-05 | 0
| 8.9 | 8.85 | 1.7713e-08 | 2.2807e-09 | 1.7702e-09 | 7 | 6.95 | 0 | 0 | 0 | 14.9 | 14.8 | 0.00056096 | 7 | 6.95 | 0.0001 |  | 0 | 0 | 9.2679e-05 | 1.413e-14
| 9 | 8.95 | 2.2e-09 | 2.659e-09 | 3.3605e-09 | 7.1 | 7.05 | 0 | 0 | 0 | 15.1 | 15 | 0.00062164 | 7.1 | 7.05 | 0.0003 |  | 0 | 0 | 0.0001027 | 4.2391e-14
| 9.1 | 9.05 | 1.4045e-08 | 1.1692e-08 | 8.1442e-09 | 7.2 | 7.15 | 0 | 0 | 0 | 15.3 | 15.2 | 0.00069289 | 7.2 | 7.15 | 0.0009 |  | 0 | 0 | 0.00011448 | 1.2717e-13
| 9.2 | 9.15 | 1.8721e-08 | 1.5892e-08 | 1.0844e-08 | 7.3 | 7.25 | 0 | 0 | 0 | 15.5 | 15.4 | 0.00073947 | 7.3 | 7.25 | 0.0023 |  | 0 | 0 | 0.00012217 | 3.25e-13
| 9.3 | 9.25 | 5.3304e-09 | 4.3136e-09 | 5.8299e-09 | 7.4 | 7.35 | 0 | 0 | 0 | 15.7 | 15.6 | 0.00078925 | 7.4 | 7.35 | 0.0057 |  | 0 | 0 | 0.00013039 | 8.0544e-13
| 9.4 | 9.35 | 5.0868e-09 | 4.2251e-09 | 3.4209e-09 | 7.5 | 7.45 | 0.0001 | 0.0002 | 0.0001 | 15.9 | 15.8 | 0.00085968 | 7.5 | 7.45 | 0.0539 |  | 2.2544e-14 | 3.7722e-14 | 0.00014203 | 7.6163e-12
| 9.5 | 9.45 | 3.7405e-09 | 4.684e-09 | 3.6142e-09 | 7.6 | 7.55 | 0.0001 | 0.0001 | 0.0001 | 16.1 | 16 | 0.00092783 | 7.6 | 7.55 | 0.0308 |  | 2.2544e-14 | 1.8861e-14 | 0.00015329 | 4.3522e-12
| 9.6 | 9.55 | 1.5665e-08 | 1.0062e-08 | 7.2075e-09 | 7.7 | 7.65 | 0.0003 | 0.0003 | 0.0003 | 16.3 | 16.2 | 0.00098666 | 7.7 | 7.65 | 0.0668 |  | 6.7633e-14 | 5.6583e-14 | 0.00016301 | 9.4391e-12
| 9.7 | 9.65 | 6.5622e-08 | 4.8025e-08 | 2.8666e-08 | 7.8 | 7.75 | 0.0007 | 0.0008 | 0.0007 | 16.5 | 16.4 | 0.0010685 | 7.8 | 7.75 | 0.1431 |  | 1.5781e-13 | 1.5089e-13 | 0.00017652 | 2.0221e-11
| 9.8 | 9.75 | 3.686e-08 | 3.1223e-08 | 2.9429e-07 | 7.9 | 7.85 | 0.0018 | 0.0019 | 0.0016 | 16.7 | 16.6 | 0.001119 | 7.9 | 7.85 | 0.2804 |  | 3.6071e-13 | 3.5836e-13 | 0.00018487 | 3.9622e-11
| 9.9 | 9.85 | 1.0978e-06 | 9.0606e-07 | 6.8072e-07 | 8 | 7.95 | 0.0041 | 0.0044 | 0.0037 | 16.9 | 16.8 | 0.0011936 | 8 | 7.95 | 0.5447 |  | 8.3415e-13 | 8.2988e-13 | 0.0001972 | 7.6969e-11
| 10 | 9.95 | 1.2811e-06 | 1.0621e-06 | 7.8871e-07 | 8.1 | 8.05 | 0.0094 | 0.0101 | 0.0085 | 17.1 | 17 | 0.0012518 | 8.1 | 8.05 | 1.0495 |  | 1.9163e-12 | 1.905e-12 | 0.00020681 | 1.483e-10
| 10.1 | 10.05 | 1.5803e-06 | 1.3109e-06 | 9.8136e-07 | 8.2 | 8.15 | 0.0199 | 0.0211 | 0.018 | 17.3 | 17.2 | 0.0013003 | 8.2 | 8.15 | 1.8698 |  | 4.058e-12 | 3.9797e-12 | 0.00021483 | 2.6421e-10
| 10.2 | 10.15 | 1.8994e-06 | 1.5904e-06 | 1.2009e-06 | 8.3 | 8.25 | 0.0414 | 0.0439 | 0.0377 | 17.5 | 17.4 | 0.0013879 | 8.3 | 8.25 | 3.3215 |  | 8.4993e-12 | 8.28e-12 | 0.0002293 | 4.6934e-10
| 10.3 | 10.25 | 1.0014e-06 | 8.2917e-07 | 6.2085e-07 | 8.4 | 8.35 | 0.0833 | 0.0881 | 0.0761 | 17.7 | 17.6 | 0.0014329 | 8.4 | 8.35 | 5.741 |  | 1.7156e-11 | 1.6617e-11 | 0.00023673 | 8.1123e-10
| 10.4 | 10.35 | 1.1506e-06 | 9.5509e-07 | 7.1548e-07 | 8.5 | 8.45 | 9.1887 | 9.7044 | 8.4185 | 17.9 | 17.8 | 0.0015201 | 8.5 | 8.45 | 561.0684 |  | 1.8979e-09 | 1.8303e-09 | 0.00025114 | 7.9282e-08
| 10.5 | 10.45 | 1.5211e-07 | 1.2655e-07 | 9.6802e-08 | 8.6 | 8.55 | 0.306 | 0.3225 | 0.2812 | 18.1 | 18 | 0.0015829 | 8.6 | 8.55 | 15.8995 |  | 6.3395e-11 | 6.0827e-11 | 0.00026152 | 2.2467e-09
| 10.6 | 10.55 | 1.6182e-07 | 1.3416e-07 | 1.0215e-07 | 8.7 | 8.65 | 0.5608 | 0.5902 | 0.5168 | 18.3 | 18.2 | 0.001623 | 8.7 | 8.65 | 25.5479 |  | 1.1651e-10 | 1.1132e-10 | 0.00026815 | 3.61e-09
| 10.7 | 10.65 | 2.1364e-07 | 1.766e-07 | 1.3647e-07 | 8.8 | 8.75 | 1.0001 | 1.0506 | 0.9241 | 18.5 | 18.4 | 0.0016907 | 8.8 | 8.75 | 40.1736 |  | 2.0833e-10 | 1.9815e-10 | 0.00027932 | 5.6767e-09
| 10.8 | 10.75 | 3.4004e-07 | 2.9005e-07 | 1.8728e-07 | 8.9 | 8.85 | 2.8424 | 2.9811 | 2.6334 | 18.7 | 18.6 | 0.0017628 | 8.9 | 8.85 | 102.2908 |  | 5.9369e-10 | 5.6226e-10 | 0.00029123 | 1.4454e-08
| 10.9 | 10.85 | 3.2538e-07 | 2.7676e-07 | 2.4248e-07 | 9 | 8.95 | 2.9456 | 3.0846 | 2.7358 | 18.9 | 18.8 | 0.0018289 | 9 | 8.95 | 93.5409 |  | 6.1677e-10 | 5.8179e-10 | 0.00030215 | 1.3218e-08
| 11 | 10.95 | 3.8322e-07 | 3.0751e-07 | 2.2831e-07 | 9.1 | 9.05 | 4.8789 | 5.1015 | 4.5422 | 19.1 | 19 | 0.0018908 | 9.1 | 9.05 | 138.8085 |  | 1.024e-09 | 9.6219e-10 | 0.00031238 | 1.9614e-08
| 11.1 | 11.05 | 5.262e-07 | 4.3802e-07 | 3.2179e-07 | 9.2 | 9.15 | 7.9124 | 8.2615 | 7.3832 | 19.3 | 19.2 | 0.0019466 | 9.2 | 9.15 | 202.586 |  | 1.6645e-09 | 1.5582e-09 | 0.0003216 | 2.8626e-08
| 11.2 | 11.15 | 6.536e-07 | 5.4231e-07 | 4.1519e-07 | 9.3 | 9.25 | 12.5774 | 13.1144 | 11.762 | 19.5 | 19.4 | 0.0019587 | 9.3 | 9.25 | 291.0447 |  | 2.6517e-09 | 2.4735e-09 | 0.0003236 | 4.1126e-08
| 11.3 | 11.25 | 8.1134e-07 | 6.5704e-07 | 4.7858e-07 | 9.4 | 9.35 | 20.0191 | 20.8466 | 18.7606 | 19.7 | 19.6 | 0.0020306 | 9.4 | 9.35 | 420.6975 |  | 4.2295e-09 | 3.9319e-09 | 0.00033548 | 5.9447e-08
| 11.4 | 11.35 | 1.1171e-06 | 9.3195e-07 | 6.6797e-07 | 9.5 | 9.45 | 29.974 | 31.1738 | 28.1463 | 19.9 | 19.8 | 0.0020624 | 9.5 | 9.45 | 573.7294 |  | 6.3454e-09 | 5.8797e-09 | 0.00034073 | 8.1071e-08
| 11.5 | 11.45 | 5.1139e-06 | 4.1892e-06 | 3.2477e-06 | 9.6 | 9.55 | 45.0284 | 46.7744 | 42.3645 | 20.1 | 20 | 0.0021861 | 9.6 | 9.55 | 788.4767 |  | 9.5509e-09 | 8.8221e-09 | 0.00036117 | 1.1142e-07
| 11.6 | 11.55 | 5.657e-06 | 4.7075e-06 | 3.4744e-06 | 9.7 | 9.65 | 172.4813 | 178.9625 | 162.578 | 20.3 | 20.2 | 0.0021774 | 9.7 | 9.65 | 2822.6413 |  | 3.6652e-08 | 3.3754e-08 | 0.00035974 | 3.9885e-07
| 11.7 | 11.65 | 7.6833e-06 | 6.3613e-06 | 4.7382e-06 | 9.8 | 9.75 | 2626.4889 | 2722.1625 | 2480.0885 | 20.5 | 20.4 | 0.0021922 | 9.8 | 9.75 | 39859.0293 |  | 5.5912e-07 | 5.1343e-07 | 0.00036218 | 5.6323e-06
| 11.8 | 11.75 | 8.3891e-06 | 6.9686e-06 | 5.2153e-06 | 9.9 | 9.85 | 140.3527 | 145.3105 | 132.7558 | 20.7 | 20.6 | 0.0022213 | 9.9 | 9.85 | 1922.3526 |  | 2.9929e-08 | 2.7407e-08 | 0.00036699 | 2.7164e-07
| 11.9 | 11.85 | 5.1245e-06 | 4.306e-06 | 3.165e-06 | 10 | 9.95 | 599.5486 | 620.0926 | 568.0268 | 20.9 | 20.8 | 0.0023423 | 10 | 9.95 | 7759.3495 |  | 1.2806e-07 | 1.1696e-07 | 0.00038698 | 1.0964e-06
| 12 | 11.95 | 5.7405e-06 | 4.936e-06 | 3.679e-06 | 10.1 | 10.05 | 3564.618 | 3683.1422 | 3382.5279 | 21.1 | 21 | 0.0023598 | 10.1 | 10.05 | 42943.1721 |  | 7.6257e-07 | 6.9468e-07 | 0.00038987 | 6.0681e-06
| 12.1 | 12.05 | 3.9455e-06 | 3.3079e-06 | 2.4413e-06 | 10.2 | 10.15 | 373.1468 | 385.1895 | 354.6228 | 21.3 | 21.2 | 0.0024007 | 10.2 | 10.15 | 4108.581 |  | 7.9948e-08 | 7.2651e-08 | 0.00039663 | 5.8056e-07
| 12.2 | 12.15 | 4.8219e-06 | 3.9534e-06 | 2.9995e-06 | 10.3 | 10.25 | 773.3436 | 797.5749 | 736.0271 | 21.5 | 21.4 | 0.002483 | 10.3 | 10.25 | 8020.7879 |  | 1.6593e-07 | 1.5043e-07 | 0.00041023 | 1.1334e-06
| 12.3 | 12.25 | 6.0475e-06 | 5.0628e-06 | 3.757e-06 | 10.4 | 10.35 | 549.3866 | 566.1042 | 523.612 | 21.7 | 21.6 | 0.0024538 | 10.4 | 10.35 | 5214.8833 |  | 1.1805e-07 | 1.0677e-07 | 0.00040539 | 7.3689e-07
| 12.4 | 12.35 | 5.998e-06 | 5.0229e-06 | 3.7907e-06 | 10.5 | 10.45 | 628.8718 | 647.4617 | 600.1789 | 21.9 | 21.8 | 0.0025212 | 10.5 | 10.45 | 5583.0684 |  | 1.3531e-07 | 1.2212e-07 | 0.00041654 | 7.8891e-07
| 12.5 | 12.45 | 7.5588e-06 | 6.24e-06 | 4.6561e-06 | 10.6 | 10.55 | 830.0544 | 853.8979 | 793.2142 | 22.1 | 22 | 0.0025031 | 10.6 | 10.55 | 6942.1051 |  | 1.7883e-07 | 1.6105e-07 | 0.00041355 | 9.8095e-07
| 12.6 | 12.55 | 7.9658e-06 | 6.5848e-06 | 4.8908e-06 | 10.7 | 10.65 | 1084.2441 | 1114.5175 | 1037.4216 | 22.3 | 22.2 | 0.0025615 | 10.7 | 10.65 | 8561.5993 |  | 2.3388e-07 | 2.1021e-07 | 0.00042319 | 1.2098e-06
| 12.7 | 12.65 | 9.3638e-06 | 7.7406e-06 | 5.7435e-06 | 10.8 | 10.75 | 1402.6874 | 1440.7665 | 1343.7346 | 22.5 | 22.4 | 0.0025877 | 10.8 | 10.75 | 10480.0268 |  | 3.0294e-07 | 2.7174e-07 | 0.00042752 | 1.4809e-06
| 12.8 | 12.75 | 1.0298e-05 | 8.5034e-06 | 6.2871e-06 | 10.9 | 10.85 | 1797.9599 | 1845.4295 | 1724.3997 | 22.7 | 22.6 | 0.0025775 | 10.9 | 10.85 | 12736.2676 |  | 3.8876e-07 | 3.4807e-07 | 0.00042584 | 1.7997e-06
| 12.9 | 12.85 | 1.1933e-05 | 9.9083e-06 | 7.3963e-06 | 11 | 10.95 | 2283.6163 | 2342.2686 | 2192.6445 | 22.9 | 22.8 | 0.0026189 | 11 | 10.95 | 15366.8817 |  | 4.9432e-07 | 4.4177e-07 | 0.00043268 | 2.1714e-06
| 13 | 12.95 | 1.2971e-05 | 1.0715e-05 | 8.0708e-06 | 11.1 | 11.05 | 3457.6811 | 3544.0958 | 3323.5315 | 23.1 | 23 | 0.0026477 | 11.1 | 11.05 | 22235.7525 |  | 7.4927e-07 | 6.6845e-07 | 0.00043743 | 3.142e-06
| 13.1 | 13.05 | 1.5658e-05 | 1.2991e-05 | 9.7427e-06 | 11.2 | 11.15 | 3590.1616 | 3677.4902 | 3454.4786 | 23.3 | 23.2 | 0.0026315 | 11.2 | 11.15 | 21919.4582 |  | 7.7879e-07 | 6.9361e-07 | 0.00043477 | 3.0973e-06
| 13.2 | 13.15 | 1.6245e-05 | 1.3783e-05 | 1.0348e-05 | 11.3 | 11.25 | 4478.9036 | 4584.9636 | 4313.9827 | 23.5 | 23.4 | 0.0026612 | 11.3 | 11.25 | 26117.1772 |  | 9.7256e-07 | 8.6477e-07 | 0.00043967 | 3.6905e-06
| 13.3 | 13.25 | 1.9773e-05 | 1.6288e-05 | 1.2056e-05 | 11.4 | 11.35 | 5467.815 | 5593.8928 | 5271.6121 | 23.7 | 23.6 | 0.0026814 | 11.4 | 11.35 | 30490.3535 |  | 1.1885e-06 | 1.0551e-06 | 0.00044301 | 4.3084e-06
| 13.4 | 13.35 | 1.9641e-05 | 1.6264e-05 | 1.2482e-05 | 11.5 | 11.45 | 45462.8408 | 46483.8563 | 43872.7215 | 23.9 | 23.8 | 0.0026984 | 11.5 | 11.45 | 247764.5804 |  | 9.8909e-06 | 8.7673e-06 | 0.00044581 | 3.501e-05
| 13.5 | 13.45 | 2.2917e-05 | 1.9105e-05 | 1.4275e-05 | 11.6 | 11.55 | 8418.3296 | 8602.5155 | 8131.2694 | 24.1 | 24 | 0.0027324 | 11.6 | 11.55 | 43182.9049 |  | 1.8332e-06 | 1.6225e-06 | 0.00045144 | 6.1019e-06
| 13.6 | 13.55 | 2.4788e-05 | 2.0549e-05 | 1.5812e-05 | 11.7 | 11.65 | 14747.9236 | 15062.3488 | 14257.5321 | 24.3 | 24.2 | 0.0027624 | 11.7 | 11.65 | 72936.5979 |  | 3.2143e-06 | 2.8409e-06 | 0.00045638 | 1.0306e-05
| 13.7 | 13.65 | 2.8869e-05 | 2.4015e-05 | 1.7996e-05 | 11.8 | 11.75 | 16811.5077 | 17160.8485 | 16266.2858 | 24.5 | 24.4 | 0.0027393 | 11.8 | 11.75 | 79852.899 |  | 3.6671e-06 | 3.2367e-06 | 0.00045257 | 1.1284e-05
| 13.8 | 13.75 | 3.038e-05 | 2.5251e-05 | 1.8881e-05 | 11.9 | 11.85 | 11936.0414 | 12177.8419 | 11558.4087 | 24.7 | 24.6 | 0.002786 | 11.9 | 11.85 | 54016.3589 |  | 2.6058e-06 | 2.2969e-06 | 0.00046028 | 7.6328e-06
| 13.9 | 13.85 | 3.332e-05 | 2.7609e-05 | 2.0634e-05 | 12 | 11.95 | 14446.4512 | 14731.8215 | 14000.487 | 24.9 | 24.8 | 0.0027406 | 12 | 11.95 | 63095.7025 |  | 3.1563e-06 | 2.7786e-06 | 0.00045279 | 8.9157e-06
| 14 | 13.95 | 3.5455e-05 | 2.9639e-05 | 2.2374e-05 | 12.1 | 12.05 | 16637.5479 | 16958.0878 | 16136.3107 | 25.1 | 25 | 0.0027694 | 12.1 | 12.05 | 70131.2934 |  | 3.6378e-06 | 3.1985e-06 | 0.00045754 | 9.9099e-06
| 14.1 | 14.05 | 4.0067e-05 | 3.329e-05 | 2.4822e-05 | 12.2 | 12.15 | 23150.4555 | 23585.557 | 22469.6661 | 25.3 | 25.2 | 0.0027973 | 12.2 | 12.15 | 94746.1854 |  | 5.0657e-06 | 4.4485e-06 | 0.00046215 | 1.3388e-05
| 14.2 | 14.15 | 4.1616e-05 | 3.4524e-05 | 2.5759e-05 | 12.3 | 12.25 | 21195.1627 | 21583.8474 | 20586.6464 | 25.5 | 25.4 | 0.0027766 | 12.3 | 12.25 | 83299.8919 |  | 4.6412e-06 | 4.0709e-06 | 0.00045873 | 1.1771e-05
| 14.3 | 14.25 | 4.6133e-05 | 3.8313e-05 | 2.8859e-05 | 12.4 | 12.35 | 24410.1733 | 24847.042 | 23725.8358 | 25.7 | 25.6 | 0.0027611 | 12.4 | 12.35 | 92918.2302 |  | 5.3489e-06 | 4.6864e-06 | 0.00045616 | 1.313e-05
| 14.4 | 14.35 | 4.8354e-05 | 4.0201e-05 | 2.9926e-05 | 12.5 | 12.45 | 28167.6039 | 28659.686 | 27396.3561 | 25.9 | 25.8 | 0.0027612 | 12.5 | 12.45 | 103991.0666 |  | 6.1764e-06 | 5.4055e-06 | 0.00045618 | 1.4694e-05
| 14.5 | 14.45 | 5.2849e-05 | 4.4434e-05 | 3.334e-05 | 12.6 | 12.55 | 32355.3287 | 32907.1874 | 31489.9352 | 26.1 | 26 | 0.0027946 | 12.6 | 12.55 | 115967.7963 |  | 7.0992e-06 | 6.2066e-06 | 0.0004617 | 1.6387e-05
| 14.6 | 14.55 | 5.5922e-05 | 4.676e-05 | 3.5006e-05 | 12.7 | 12.65 | 36995.5987 | 37611.7829 | 36028.8389 | 26.3 | 26.2 | 0.0027628 | 12.7 | 12.65 | 128852.2956 |  | 8.1225e-06 | 7.094e-06 | 0.00045646 | 1.8207e-05
| 14.7 | 14.65 | 6.2283e-05 | 5.1495e-05 | 3.8909e-05 | 12.8 | 12.75 | 42115.4593 | 42800.5736 | 41040.0184 | 26.5 | 26.4 | 0.0028174 | 12.8 | 12.75 | 142667.1828 |  | 9.2523e-06 | 8.0726e-06 | 0.00046548 | 2.016e-05
| 14.8 | 14.75 | 6.3409e-05 | 5.2878e-05 | 3.9932e-05 | 12.9 | 12.85 | 47745.4703 | 48504.2177 | 46553.8729 | 26.7 | 26.6 | 0.0027895 | 12.9 | 12.85 | 157445.8354 |  | 1.0495e-05 | 9.1484e-06 | 0.00046086 | 2.2248e-05
| 14.9 | 14.85 | 7.0065e-05 | 5.8677e-05 | 4.3863e-05 | 13 | 12.95 | 53912.5203 | 54749.6242 | 52597.2525 | 26.9 | 26.8 | 0.0027386 | 13 | 12.95 | 173207.4563 |  | 1.1858e-05 | 1.0326e-05 | 0.00045245 | 2.4475e-05
| 15 | 14.95 | 7.2996e-05 | 6.0619e-05 | 4.5219e-05 | 13.1 | 13.05 | 60642.5401 | 61562.7178 | 59196.0919 | 27.1 | 27 | 0.0028031 | 13.1 | 13.05 | 189968.3995 |  | 1.3345e-05 | 1.1611e-05 | 0.0004631 | 2.6843e-05
| 15.1 | 15.05 | 7.9813e-05 | 6.6914e-05 | 5.066e-05 | 13.2 | 13.15 | 67960.7278 | 68968.6697 | 66375.626 | 27.3 | 27.2 | 0.0027389 | 13.2 | 13.15 | 207741.2317 |  | 1.4964e-05 | 1.3008e-05 | 0.0004525 | 2.9355e-05
| 15.2 | 15.15 | 8.0745e-05 | 6.7341e-05 | 5.061e-05 | 13.3 | 13.25 | 75891.1458 | 76991.4913 | 74159.9924 | 27.5 | 27.4 | 0.0026887 | 13.3 | 13.25 | 226536.4142 |  | 1.6719e-05 | 1.4521e-05 | 0.00044421 | 3.2011e-05
| 15.3 | 15.25 | 8.8105e-05 | 7.4253e-05 | 5.5275e-05 | 13.4 | 13.35 | 84456.3335 | 85653.6441 | 82571.8488 | 27.7 | 27.6 | 0.0026991 | 13.4 | 13.35 | 246361.5666 |  | 1.8615e-05 | 1.6155e-05 | 0.00044593 | 3.4812e-05
| 15.4 | 15.35 | 9.0418e-05 | 7.5275e-05 | 5.7031e-05 | 13.5 | 13.45 | 93677.9466 | 94976.6904 | 91632.9933 | 27.9 | 27.8 | 0.0027593 | 13.5 | 13.45 | 267220.7176 |  | 2.0658e-05 | 1.7914e-05 | 0.00045587 | 3.776e-05
| 15.5 | 15.45 | 9.8105e-05 | 8.1513e-05 | 6.1313e-05 | 13.6 | 13.55 | 103575.4852 | 104980.0067 | 101363.1174 | 28.1 | 28 | 0.0027314 | 13.6 | 13.55 | 289115.0287 |  | 2.2852e-05 | 1.98e-05 | 0.00045127 | 4.0853e-05
| 15.6 | 15.55 | 0.00010248 | 8.5674e-05 | 6.4654e-05 | 13.7 | 13.65 | 114166.8809 | 115681.3815 | 111780.3747 | 28.3 | 28.2 | 0.0026985 | 13.7 | 13.65 | 312041.5737 |  | 2.52e-05 | 2.1819e-05 | 0.00044582 | 4.4093e-05
| 15.7 | 15.65 | 0.00010796 | 9.0082e-05 | 6.7427e-05 | 13.8 | 13.75 | 125468.4492 | 127096.9699 | 122901.3313 | 28.5 | 28.4 | 0.0027085 | 13.8 | 13.75 | 335994.3868 |  | 2.7707e-05 | 2.3972e-05 | 0.00044748 | 4.7478e-05
| 15.8 | 15.75 | 0.00011172 | 9.3542e-05 | 7.0426e-05 | 13.9 | 13.85 | 137494.8365 | 139241.2417 | 134740.9093 | 28.7 | 28.6 | 0.0026648 | 13.9 | 13.85 | 360965.933 |  | 3.0377e-05 | 2.6262e-05 | 0.00044026 | 5.1006e-05
| 15.9 | 15.85 | 0.00012041 | 0.00010011 | 7.4531e-05 | 14 | 13.95 | 150259.7687 | 152127.7417 | 147313.1188 | 28.9 | 28.8 | 0.002658 | 14 | 13.95 | 386944.6474 |  | 3.3211e-05 | 2.8693e-05 | 0.00043914 | 5.4677e-05
| 16 | 15.95 | 0.00012544 | 0.00010435 | 7.8733e-05 | 14.1 | 14.05 | 163771.4593 | 165764.4427 | 160626.5515 | 29.1 | 29 | 0.0026555 | 14.1 | 14.05 | 413917.8877 |  | 3.6212e-05 | 3.1265e-05 | 0.00043873 | 5.8489e-05
| 16.1 | 16.05 | 0.00013233 | 0.00011084 | 8.3953e-05 | 14.2 | 14.15 | 178016.786 | 180137.7781 | 174668.7865 | 29.3 | 29.2 | 0.0025781 | 14.2 | 14.15 | 441821.6387 |  | 3.9378e-05 | 3.3976e-05 | 0.00042594 | 6.2431e-05
| 16.2 | 16.15 | 0.00013698 | 0.00011388 | 8.5221e-05 | 14.3 | 14.25 | 192998.2014 | 195249.9399 | 189442.6878 | 29.5 | 29.4 | 0.0026682 | 14.3 | 14.25 | 470628.7941 |  | 4.2709e-05 | 3.6826e-05 | 0.00044083 | 6.6502e-05
| 16.3 | 16.25 | 0.00014484 | 0.00012141 | 9.139e-05 | 14.4 | 14.35 | 208741.5588 | 211126.7625 | 204974.1379 | 29.7 | 29.6 | 0.0026481 | 14.4 | 14.35 | 500363.8928 |  | 4.621e-05 | 3.9821e-05 | 0.00043751 | 7.0704e-05
| 16.4 | 16.35 | 0.00014947 | 0.00012429 | 9.3663e-05 | 14.5 | 14.45 | 225249.6333 | 227770.7726 | 221266.3049 | 29.9 | 29.8 | 0.0026153 | 14.5 | 14.45 | 531003.937 |  | 4.9883e-05 | 4.296e-05 | 0.00043208 | 7.5033e-05
| 16.5 | 16.45 | 0.00015749 | 0.00013219 | 9.9804e-05 | 14.6 | 14.55 | 242519.5224 | 245178.7716 | 238316.7547 | 30.1 | 30 | 0.0025868 | 14.6 | 14.55 | 562527.1179 |  | 5.3727e-05 | 4.6243e-05 | 0.00042737 | 7.9488e-05
| 16.6 | 16.55 | 0.00016236 | 0.00013648 | 0.00010229 | 14.7 | 14.65 | 260550.4435 | 263349.7229 | 256125.1075 | 30.3 | 30.2 | 0.0026008 | 14.7 | 14.65 | 594903.0062 |  | 5.7742e-05 | 4.967e-05 | 0.00042969 | 8.4063e-05
| 16.7 | 16.65 | 0.00017062 | 0.00014296 | 0.00010798 | 14.8 | 14.75 | 279338.78 | 282279.7418 | 274688.1719 | 30.5 | 30.4 | 0.0025402 | 14.8 | 14.75 | 628096.1444 |  | 6.1927e-05 | 5.3241e-05 | 0.00041967 | 8.8753e-05
| 16.8 | 16.75 | 0.00017364 | 0.00014548 | 0.00011054 | 14.9 | 14.85 | 298880.3797 | 301964.4199 | 294002.2028 | 30.7 | 30.6 | 0.0025558 | 14.9 | 14.85 | 662080.4801 |  | 6.6281e-05 | 5.6953e-05 | 0.00042225 | 9.3555e-05
| 16.9 | 16.85 | 0.0001843 | 0.00015359 | 0.00011553 | 15 | 14.95 | 319166.9113 | 322395.1443 | 314059.3163 | 30.9 | 30.8 | 0.0025652 | 15 | 14.95 | 696822.4309 |  | 7.0803e-05 | 6.0807e-05 | 0.0004238 | 9.8464e-05
| 17 | 16.95 | 0.00018643 | 0.00015616 | 0.00011814 | 15.1 | 15.05 | 340184.245 | 343557.4651 | 334845.8912 | 31.1 | 31 | 0.0025076 | 15.1 | 15.05 | 732294.438 |  | 7.5489e-05 | 6.4798e-05 | 0.00041429 | 0.00010348
| 17.1 | 17.05 | 0.00020038 | 0.0001679 | 0.00012644 | 15.2 | 15.15 | 361857.3038 | 365375.5777 | 356288.0084 | 31.3 | 31.2 | 0.0025553 | 15.2 | 15.15 | 768342.2589 |  | 8.0323e-05 | 6.8913e-05 | 0.00042216 | 0.00010857
| 17.2 | 17.15 | 0.00020909 | 0.00017368 | 0.00013064 | 15.3 | 15.25 | 384167.0636 | 387830.1874 | 378367.0751 | 31.5 | 31.4 | 0.0025179 | 15.3 | 15.25 | 804917.0289 |  | 8.5301e-05 | 7.3149e-05 | 0.00041599 | 0.00011374
| 17.3 | 17.25 | 0.00021306 | 0.00017776 | 0.00013495 | 15.4 | 15.35 | 407162.7044 | 410970.6879 | 401131.9336 | 31.7 | 31.6 | 0.0024261 | 15.4 | 15.35 | 842096.3112 |  | 9.0433e-05 | 7.7513e-05 | 0.00040083 | 0.00011899
| 17.4 | 17.35 | 0.00021556 | 0.00018115 | 0.00013816 | 15.5 | 15.45 | 430825.5236 | 434778.0955 | 424564.3295 | 31.9 | 31.8 | 0.0024488 | 15.5 | 15.45 | 879846.0413 |  | 9.5716e-05 | 8.2003e-05 | 0.00040457 | 0.00012433
| 17.5 | 17.45 | 0.00022595 | 0.00018879 | 0.00014199 | 15.6 | 15.55 | 455135.4189 | 459232.0343 | 448644.5967 | 32.1 | 32 | 0.0023572 | 15.6 | 15.55 | 918134.0256 |  | 0.00010114 | 8.6616e-05 | 0.00038945 | 0.00012974
| 17.6 | 17.55 | 0.00023251 | 0.00019439 | 0.00014669 | 15.7 | 15.65 | 480077.3418 | 484317.2474 | 473358.0208 | 32.3 | 32.2 | 0.002376 | 15.7 | 15.65 | 956912.9125 |  | 0.00010672 | 9.1347e-05 | 0.00039255 | 0.00013522
| 17.7 | 17.65 | 0.00024418 | 0.00020436 | 0.00015446 | 15.8 | 15.75 | 505631.3552 | 510013.5622 | 498685.0411 | 32.5 | 32.4 | 0.0023445 | 15.8 | 15.75 | 996137.9517 |  | 0.00011243 | 9.6194e-05 | 0.00038734 | 0.00014076
| 17.8 | 17.75 | 0.00024724 | 0.00020738 | 0.00015725 | 15.9 | 15.85 | 531779.0003 | 536302.3148 | 524607.5276 | 32.7 | 32.6 | 0.0024062 | 15.9 | 15.85 | 1035777.0603 |  | 0.00011827 | 0.00010115 | 0.00039754 | 0.00014636
| 17.9 | 17.85 | 0.00025622 | 0.00021578 | 0.00016369 | 16 | 15.95 | 558489.3635 | 563152.2981 | 551095.0354 | 32.9 | 32.8 | 0.0023564 | 16 | 15.95 | 1075796.6622 |  | 0.00012424 | 0.00010622 | 0.00038932 | 0.00015202
| 18 | 17.95 | 0.00026233 | 0.00022053 | 0.00016721 | 16.1 | 16.05 | 585742.3758 | 590543.261 | 578127.7881 | 33.1 | 33 | 0.0023076 | 16.1 | 16.05 | 1116158.7641 |  | 0.00013034 | 0.00011138 | 0.00038124 | 0.00015772
| 18.1 | 18.05 | 0.000271 | 0.00022619 | 0.0001709 | 16.2 | 16.15 | 613513.8404 | 618450.8059 | 605681.9113 | 33.3 | 33.2 | 0.0023102 | 16.2 | 16.15 | 1156818.8055 |  | 0.00013655 | 0.00011665 | 0.00038168 | 0.00016346
| 18.2 | 18.15 | 0.00027817 | 0.00023248 | 0.00017554 | 16.3 | 16.25 | 641782.3691 | 646853.3818 | 633736.2785 | 33.5 | 33.4 | 0.0023068 | 16.3 | 16.25 | 1197729.482 |  | 0.00014287 | 0.000122 | 0.00038112 | 0.00016924
| 18.3 | 18.25 | 0.00028564 | 0.00023921 | 0.0001812 | 16.4 | 16.35 | 670478.8562 | 675681.3634 | 662222.612 | 33.7 | 33.6 | 0.0022842 | 16.4 | 16.35 | 1238776.3445 |  | 0.00014929 | 0.00012744 | 0.00037738 | 0.00017504
| 18.4 | 18.35 | 0.00028733 | 0.00024101 | 0.00018222 | 16.5 | 16.45 | 699553.1535 | 704884.2666 | 691091.3011 | 33.9 | 33.8 | 0.0022691 | 16.5 | 16.45 | 1279912.9117 |  | 0.0001558 | 0.00013295 | 0.00037489 | 0.00018086
| 18.5 | 18.45 | 0.00029704 | 0.00024848 | 0.00018817 | 16.6 | 16.55 | 729039.3984 | 734496.5501 | 720375.977 | 34.1 | 34 | 0.0022063 | 16.6 | 16.55 | 1321198.6304 |  | 0.0001624 | 0.00013853 | 0.0003645 | 0.00018669
| 18.6 | 18.55 | 0.00029978 | 0.00025096 | 0.0001889 | 16.7 | 16.65 | 758911.1667 | 764491.6459 | 750050.4471 | 34.3 | 34.2 | 0.0021866 | 16.7 | 16.65 | 1362588.1877 |  | 0.00016909 | 0.00014419 | 0.00036125 | 0.00019254
| 18.7 | 18.65 | 0.0003081 | 0.00025831 | 0.00019676 | 16.8 | 16.75 | 789143.3549 | 794844.329 | 780089.8047 | 34.5 | 34.4 | 0.0021205 | 16.8 | 16.75 | 1404037.9642 |  | 0.00017587 | 0.00014992 | 0.00035033 | 0.0001984
| 18.8 | 18.75 | 0.00031552 | 0.00026383 | 0.0001995 | 16.9 | 16.85 | 819711.5991 | 825530.1305 | 810469.8562 | 34.7 | 34.6 | 0.0021517 | 16.9 | 16.85 | 1445521.2528 |  | 0.00018272 | 0.0001557 | 0.00035549 | 0.00020426
| 18.9 | 18.85 | 0.00032371 | 0.00027063 | 0.00020395 | 17 | 16.95 | 850572.3488 | 856505.2709 | 841147.4181 | 34.9 | 34.8 | 0.0021535 | 17 | 16.95 | 1487007.5031 |  | 0.00018963 | 0.00016155 | 0.00035578 | 0.00021012
| 19 | 18.95 | 0.00032499 | 0.00027258 | 0.00020852 | 17.1 | 17.05 | 881708.608 | 887752.7261 | 872105.5424 | 35.1 | 35 | 0.0021051 | 17.1 | 17.05 | 1528474.6256 |  | 0.00019661 | 0.00016744 | 0.00034779 | 0.00021598
| 19.1 | 19.05 | 0.00033236 | 0.00027929 | 0.00021146 | 17.2 | 17.15 | 913092.1292 | 919244.15 | 903316.1418 | 35.3 | 35.2 | 0.0021011 | 17.2 | 17.15 | 1569868.8577 |  | 0.00020365 | 0.00017338 | 0.00034712 | 0.00022183
| 19.2 | 19.15 | 0.00033914 | 0.00028355 | 0.00021461 | 17.3 | 17.25 | 944703.8698 | 950960.472 | 934760.2213 | 35.5 | 35.4 | 0.0019842 | 17.3 | 17.25 | 1611156.9122 |  | 0.00021074 | 0.00017936 | 0.00032782 | 0.00022766
| 19.3 | 19.25 | 0.00034589 | 0.00029191 | 0.00022023 | 17.4 | 17.35 | 976506.5468 | 982864.2674 | 966400.7258 | 35.7 | 35.6 | 0.0020534 | 17.4 | 17.35 | 1652320.1648 |  | 0.00021787 | 0.00018538 | 0.00033924 | 0.00023348
| 19.4 | 19.35 | 0.0003474 | 0.00029124 | 0.00022056 | 17.5 | 17.45 | 1008466.4339 | 1014921.7045 | 998204.1006 | 35.9 | 35.8 | 0.0020218 | 17.5 | 17.45 | 1693332.0421 |  | 0.00022504 | 0.00019142 | 0.00033402 | 0.00023928
| 19.5 | 19.45 | 0.00036537 | 0.00030638 | 0.00023301 | 17.6 | 17.55 | 1040565.2501 | 1047114.5037 | 1030152.0645 | 36.1 | 36 | 0.0019568 | 17.6 | 17.55 | 1734175.3824 |  | 0.00023224 | 0.0001975 | 0.00032329 | 0.00024505
| 19.6 | 19.55 | 0.00037086 | 0.0003096 | 0.00023371 | 17.7 | 17.65 | 1072508.9151 | 1079147.5466 | 1061952.1996 | 36.3 | 36.2 | 0.0019614 | 17.7 | 17.65 | 1774431.084 |  | 0.00023941 | 0.00020354 | 0.00032405 | 0.00025074
| 19.7 | 19.65 | 0.0003791 | 0.00031876 | 0.00024209 | 17.8 | 17.75 | 1104269.1457 | 1110992.6016 | 1093576.1437 | 36.5 | 36.4 | 0.0019525 | 17.8 | 17.75 | 1814067.6417 |  | 0.00024654 | 0.00020954 | 0.00032258 | 0.00025634
| 19.8 | 19.75 | 0.0003867 | 0.00032403 | 0.0002446 | 17.9 | 17.85 | 1136068.7786 | 1142873.3853 | 1125245.3289 | 36.7 | 36.6 | 0.0019455 | 17.9 | 17.85 | 1853432.1262 |  | 0.00025368 | 0.00021556 | 0.00032142 | 0.0002619
| 19.9 | 19.85 | 0.00038723 | 0.00032564 | 0.00024751 | 18 | 17.95 | 1167886.3729 | 1174768.4592 | 1156938.3129 | 36.9 | 36.8 | 0.0018438 | 18 | 17.95 | 1892503.2583 |  | 0.00026083 | 0.00022157 | 0.00030462 | 0.00026742
| 20 | 19.95 | 0.00039179 | 0.00032847 | 0.00024773 | 18.1 | 18.05 | 1199705.3659 | 1206661.2947 | 1188638.4811 | 37.1 | 37 | 0.0018048 | 18.1 | 18.05 | 1931277.9137 |  | 0.00026797 | 0.00022759 | 0.00029818 | 0.0002729
| 20.1 | 20.05 | 0.00039225 | 0.00032958 | 0.00025114 | 18.2 | 18.15 | 1231505.1506 | 1238531.297 | 1220325.2093 | 37.3 | 37.2 | 0.0018166 | 18.2 | 18.15 | 1969702.4351 |  | 0.00027512 | 0.0002336 | 0.00030012 | 0.00027833
| 20.2 | 20.15 | 0.00039285 | 0.00033074 | 0.00024942 | 18.3 | 18.25 | 1263275.9822 | 1270368.7971 | 1251988.6343 | 37.5 | 37.4 | 0.0017292 | 18.3 | 18.25 | 2007751.763 |  | 0.00028225 | 0.0002396 | 0.00028569 | 0.0002837
| 20.3 | 20.25 | 0.0003997 | 0.00033491 | 0.00025515 | 18.4 | 18.35 | 1294955.5832 | 1302111.3015 | 1283566.8244 | 37.7 | 37.6 | 0.0017226 | 18.4 | 18.35 | 2045423.1872 |  | 0.00028937 | 0.00024559 | 0.00028459 | 0.00028903
| 20.4 | 20.35 | 0.00040351 | 0.00033885 | 0.00025907 | 18.5 | 18.45 | 1326542.8387 | 1333757.8289 | 1315058.4541 | 37.9 | 37.8 | 0.0016908 | 18.5 | 18.45 | 2082704.3505 |  | 0.00029647 | 0.00025156 | 0.00027935 | 0.0002943
| 20.5 | 20.45 | 0.00041063 | 0.00034456 | 0.0002619 | 18.6 | 18.55 | 1358023.8329 | 1365294.5236 | 1346449.514 | 38.1 | 38 | 0.0016747 | 18.6 | 18.55 | 2119589.6577 |  | 0.00030355 | 0.00025751 | 0.00027668 | 0.00029951
| 20.6 | 20.55 | 0.00041551 | 0.00034863 | 0.00026457 | 18.7 | 18.65 | 1389377.8887 | 1396700.7322 | 1377719.2904 | 38.3 | 38.2 | 0.001692 | 18.7 | 18.65 | 2156032.933 |  | 0.0003106 | 0.00026343 | 0.00027955 | 0.00030466
| 20.7 | 20.65 | 0.00042063 | 0.00035353 | 0.00026901 | 18.8 | 18.75 | 1420591.6913 | 1427963.2042 | 1408854.3687 | 38.5 | 38.4 | 0.0016085 | 18.8 | 18.75 | 2192009.8945 |  | 0.00031762 | 0.00026933 | 0.00026575 | 0.00030974
| 20.8 | 20.75 | 0.00042651 | 0.00035858 | 0.00027281 | 18.9 | 18.85 | 1451614.8656 | 1459031.4378 | 1439804.5765 | 38.7 | 38.6 | 0.0016612 | 18.9 | 18.85 | 2227523.6713 |  | 0.0003246 | 0.00027519 | 0.00027445 | 0.00031476
| 20.9 | 20.85 | 0.0004296 | 0.00036155 | 0.00027503 | 19 | 18.95 | 1482449.2268 | 1489907.3957 | 1470571.4962 | 38.9 | 38.8 | 0.0015705 | 19 | 18.95 | 2262566.8748 |  | 0.00033153 | 0.00028101 | 0.00025947 | 0.00031971
| 21 | 20.95 | 0.00043563 | 0.00036588 | 0.00027936 | 19.1 | 19.05 | 1513076.2537 | 1520572.597 | 1501136.5436 | 39.1 | 39 | 0.0015615 | 19.1 | 19.05 | 2297142.9783 |  | 0.00033842 | 0.00028679 | 0.00025798 | 0.0003246
| 21.1 | 21.05 | 0.00044084 | 0.00037168 | 0.00028185 | 19.2 | 19.15 | 1543494.6145 | 1551025.8348 | 1531498.189 | 39.3 | 39.2 | 0.001543 | 19.2 | 19.15 | 2331202.7257 |  | 0.00034527 | 0.00029254 | 0.00025492 | 0.00032941
| 21.2 | 21.15 | 0.00044696 | 0.00037554 | 0.00028522 | 19.3 | 19.25 | 1573671.7412 | 1581234.5115 | 1561623.9133 | 39.5 | 39.4 | 0.0014663 | 19.3 | 19.25 | 2364724.986 |  | 0.00035206 | 0.00029824 | 0.00024224 | 0.00033415
| 21.3 | 21.25 | 0.00045512 | 0.00038238 | 0.0002903 | 19.4 | 19.35 | 1603575.5137 | 1611166.484 | 1591481.6342 | 39.7 | 39.6 | 0.0015012 | 19.4 | 19.35 | 2397725.814 |  | 0.00035879 | 0.00030388 | 0.00024802 | 0.00033881
| 21.4 | 21.35 | 0.00045454 | 0.00038258 | 0.00029104 | 19.5 | 19.45 | 1633207.8393 | 1640823.7994 | 1621073.0373 | 39.9 | 39.8 | 0.0014849 | 19.5 | 19.45 | 2430203.9115 |  | 0.00036546 | 0.00030948 | 0.00024532 | 0.0003434
| 21.5 | 21.45 | 0.00046342 | 0.00038962 | 0.0002972 | 19.6 | 19.55 | 1662559.4266 | 1670197.2491 | 1650388.7001 | 40.1 | 40 | 0.0014532 | 19.6 | 19.55 | 2462158.2507 |  | 0.00037207 | 0.00031502 | 0.00024008 | 0.00034791
| 21.6 | 21.55 | 0.0004686 | 0.00039411 | 0.00030149 | 19.7 | 19.65 | 1691628.2202 | 1699284.8914 | 1679426.3873 | 40.3 | 40.2 | 0.001405 | 19.7 | 19.65 | 2493541.7278 |  | 0.00037862 | 0.0003205 | 0.00023212 | 0.00035235
| 21.7 | 21.65 | 0.00047608 | 0.0004002 | 0.00030364 | 19.8 | 19.75 | 1720373.6256 | 1728046.0692 | 1708145.6052 | 40.5 | 40.4 | 0.0013769 | 19.8 | 19.75 | 2524354.4378 |  | 0.00038509 | 0.00032593 | 0.00022748 | 0.0003567
| 21.8 | 21.75 | 0.00047801 | 0.00040073 | 0.0003045 | 19.9 | 19.85 | 1748777.4367 | 1756462.6183 | 1736528.0812 | 40.7 | 40.6 | 0.0013022 | 19.9 | 19.85 | 2554596.1745 |  | 0.00039149 | 0.00033129 | 0.00021514 | 0.00036098
| 21.9 | 21.85 | 0.00051278 | 0.00043132 | 0.00032803 | 20 | 19.95 | 1776847.5152 | 1784542.5556 | 1764581.4309 | 40.9 | 40.8 | 0.0013238 | 20 | 19.95 | 2584276.0782 |  | 0.00039782 | 0.00033658 | 0.0002187 | 0.00036517
| 22 | 21.95 | 0.00051268 | 0.00043038 | 0.00032687 | 20.1 | 20.05 | 1804576.0691 | 1812278.1708 | 1792297.733 | 41.1 | 41 | 0.0013123 | 20.1 | 20.05 | 2613406.3115 |  | 0.00040406 | 0.00034181 | 0.0002168 | 0.00036929
| 22.1 | 22.05 | 0.00053969 | 0.00045489 | 0.00034546 | 20.2 | 20.15 | 1831599.4946 | 1839305.2884 | 1819314.3026 | 41.3 | 41.2 | 0.001248 | 20.2 | 20.15 | 2641497.8749 |  | 0.00041015 | 0.00034691 | 0.00020618 | 0.00037326
| 22.2 | 22.15 | 0.00054274 | 0.00045659 | 0.00034855 | 20.3 | 20.25 | 1857863.8557 | 1865569.9158 | 1845577.2947 | 41.5 | 41.4 | 0.0011979 | 20.3 | 20.25 | 2668559.4145 |  | 0.00041608 | 0.00035186 | 0.00019791 | 0.00037708
| 22.3 | 22.25 | 0.00052992 | 0.00044622 | 0.00033984 | 20.4 | 20.35 | 1883745.4483 | 1891449.2142 | 1871461.6188 | 41.7 | 41.6 | 0.0011756 | 20.4 | 20.35 | 2695043.5123 |  | 0.00042191 | 0.00035675 | 0.00019422 | 0.00038082
| 22.4 | 22.35 | 0.00053051 | 0.00044622 | 0.00034011 | 20.5 | 20.45 | 1909249.7173 | 1916948.7535 | 1896972.5211 | 41.9 | 41.8 | 0.001194 | 20.5 | 20.45 | 2720957.8757 |  | 0.00042766 | 0.00036156 | 0.00019727 | 0.00038448
| 22.5 | 22.45 | 0.00050866 | 0.00042708 | 0.00032441 | 20.6 | 20.55 | 1934367.5623 | 1942059.4955 | 1922100.8023 | 42.1 | 42 | 0.0011641 | 20.6 | 20.55 | 2746319.394 |  | 0.00043333 | 0.00036629 | 0.00019233 | 0.00038807
| 22.6 | 22.55 | 0.00050949 | 0.00042931 | 0.00032613 | 20.7 | 20.65 | 1959120.0238 | 1966802.6593 | 1946867.2182 | 42.3 | 42.2 | 0.0011736 | 20.7 | 20.65 | 2771079.385 |  | 0.00043891 | 0.00037096 | 0.0001939 | 0.00039157
| 22.7 | 22.65 | 0.00051589 | 0.00043393 | 0.00033012 | 20.8 | 20.75 | 1983420.7617 | 1991091.6603 | 1971185.8192 | 42.5 | 42.4 | 0.0011123 | 20.8 | 20.75 | 2795228.0006 |  | 0.00044439 | 0.00037554 | 0.00018376 | 0.00039498
| 22.8 | 22.75 | 0.000518 | 0.00043614 | 0.00033329 | 20.9 | 20.85 | 2007317.6108 | 2014974.6157 | 1995103.9904 | 42.7 | 42.6 | 0.0010677 | 20.9 | 20.85 | 2818799.0001 |  | 0.00044979 | 0.00038004 | 0.00017639 | 0.00039831
| 22.9 | 22.85 | 0.00052508 | 0.00044287 | 0.00033778 | 21 | 20.95 | 2030806.4023 | 2038447.4293 | 2018617.4468 | 42.9 | 42.8 | 0.0010796 | 21 | 20.95 | 2841799.8775 |  | 0.00045509 | 0.00038447 | 0.00017837 | 0.00040156
| 23 | 22.95 | 0.00053008 | 0.00044586 | 0.00033945 | 21.1 | 21.05 | 2053879.6614 | 2061502.6848 | 2041718.6208 | 43.1 | 43 | 0.0010014 | 21.1 | 21.05 | 2864240.8625 |  | 0.00046029 | 0.00038882 | 0.00016544 | 0.00040473
| 23.1 | 23.05 | 0.00053333 | 0.00044954 | 0.00034323 | 21.2 | 21.15 | 2076540.7877 | 2084143.8783 | 2064410.7582 | 43.3 | 43.2 | 0.00098078 | 21.2 | 21.15 | 2886079.9079 |  | 0.00046541 | 0.00039309 | 0.00016204 | 0.00040782
| 23.2 | 23.15 | 0.0005385 | 0.00045243 | 0.00034416 | 21.3 | 21.25 | 2098742.9379 | 2106324.0777 | 2086647.1574 | 43.5 | 43.4 | 0.00091576 | 21.3 | 21.25 | 2907311.9053 |  | 0.00047042 | 0.00039727 | 0.0001513 | 0.00041082
| 23.3 | 23.25 | 0.00054159 | 0.00045483 | 0.00034699 | 21.4 | 21.35 | 2120513.2033 | 2128070.5553 | 2108454.6215 | 43.7 | 43.6 | 0.00090989 | 21.4 | 21.35 | 2927972.6211 |  | 0.00047534 | 0.00040138 | 0.00015033 | 0.00041374
| 23.4 | 23.35 | 0.00054514 | 0.00045959 | 0.00034869 | 21.5 | 21.45 | 2141852.9477 | 2149384.7571 | 2129834.3828 | 43.9 | 43.8 | 0.00088131 | 21.5 | 21.45 | 2948069.5889 |  | 0.00048016 | 0.0004054 | 0.0001456 | 0.00041658
| 23.5 | 23.45 | 0.00054603 | 0.00046128 | 0.00035136 | 21.6 | 21.55 | 2162765.7052 | 2170270.3039 | 2150789.8374 | 44.1 | 44 | 0.00084963 | 21.6 | 21.55 | 2967615.1739 |  | 0.00048488 | 0.00040933 | 0.00014037 | 0.00041934
| 23.6 | 23.55 | 0.00055105 | 0.00046315 | 0.00035428 | 21.7 | 21.65 | 2183232.3106 | 2190708.0361 | 2171301.811 | 44.3 | 44.2 | 0.00088152 | 21.7 | 21.65 | 2986570.3575 |  | 0.00048951 | 0.00041319 | 0.00014564 | 0.00042202
| 23.7 | 23.65 | 0.0005539 | 0.00046695 | 0.00035556 | 21.8 | 21.75 | 2203241.7362 | 2210686.96 | 2191359.2217 | 44.5 | 44.4 | 0.00085036 | 21.8 | 21.75 | 3004930.7851 |  | 0.00049403 | 0.00041696 | 0.00014049 | 0.00042461
| 23.8 | 23.75 | 0.00055922 | 0.00047079 | 0.00035904 | 21.9 | 21.85 | 2222801.5361 | 2230214.7252 | 2210969.4713 | 44.7 | 44.6 | 0.00079982 | 21.9 | 21.85 | 3022731.1785 |  | 0.00049845 | 0.00042064 | 0.00013214 | 0.00042713
| 23.9 | 23.85 | 0.00055899 | 0.00047138 | 0.00036033 | 22 | 21.95 | 2241923.5588 | 2249303.2858 | 2230144.2399 | 44.9 | 44.8 | 0.00074569 | 22 | 21.95 | 3039984.8165 |  | 0.00050277 | 0.00042424 | 0.0001232 | 0.00042956
| 24 | 23.95 | 0.00056728 | 0.00047769 | 0.00036447 | 22.1 | 22.05 | 2260636.3893 | 2267981.3825 | 2248911.8637 | 45.1 | 45 | 0.00071784 | 22.1 | 22.05 | 3056705.6817 |  | 0.00050701 | 0.00042776 | 0.0001186 | 0.00043193
| 24.1 | 24.05 | 0.00057027 | 0.00048043 | 0.00036598 | 22.2 | 22.15 | 2278440.4021 | 2285748.887 | 2266773.5266 | 45.3 | 45.2 | 0.00071367 | 22.2 | 22.15 | 3072374.6852 |  | 0.00051103 | 0.00043111 | 0.00011791 | 0.00043414
| 24.2 | 24.15 | 0.00057158 | 0.0004825 | 0.00036946 | 22.3 | 22.25 | 2295371.3398 | 2302641.748 | 2283764.642 | 45.5 | 45.4 | 0.00069546 | 22.3 | 22.25 | 3087017.2541 |  | 0.00051486 | 0.0004343 | 0.0001149 | 0.00043621
| 24.3 | 24.25 | 0.00057335 | 0.00048218 | 0.00036716 | 22.4 | 22.35 | 2311866.5878 | 2319097.7745 | 2300321.9119 | 45.7 | 45.6 | 0.00065633 | 22.4 | 22.35 | 3101129.9159 |  | 0.0005186 | 0.0004374 | 0.00010843 | 0.0004382
| 24.4 | 24.35 | 0.00057394 | 0.00048279 | 0.00036801 | 22.5 | 22.45 | 2327910.8825 | 2335101.7074 | 2316430.0656 | 45.9 | 45.8 | 0.00061947 | 22.5 | 22.45 | 3114725.7876 |  | 0.00052223 | 0.00044042 | 0.00010234 | 0.00044013
| 24.5 | 24.45 | 0.0005783 | 0.00048758 | 0.00037182 | 22.6 | 22.55 | 2343590.1733 | 2350739.8096 | 2332174.5515 | 46.1 | 46 | 0.00056863 | 22.6 | 22.55 | 3127830.8137 |  | 0.00052578 | 0.00044337 | 9.3945e-05 | 0.00044198
| 24.6 | 24.55 | 0.00058115 | 0.00048706 | 0.000371 | 22.7 | 22.65 | 2358766.9118 | 2365874.1583 | 2347418.4192 | 46.3 | 46.2 | 0.00053421 | 22.7 | 22.65 | 3140395.3648 |  | 0.00052921 | 0.00044623 | 8.8258e-05 | 0.00044375
| 24.7 | 24.65 | 0.00058735 | 0.00049493 | 0.00037694 | 22.8 | 22.75 | 2373509.217 | 2380573.1289 | 2362229.378 | 46.5 | 46.4 | 0.00052625 | 22.8 | 22.75 | 3152437.1255 |  | 0.00053255 | 0.000449 | 8.6943e-05 | 0.00044545
| 24.8 | 24.75 | 0.00058993 | 0.0004969 | 0.00037936 | 22.9 | 22.85 | 2387827.7164 | 2394847.4245 | 2376617.9344 | 46.7 | 46.6 | 0.00045187 | 22.9 | 22.85 | 3163968.3094 |  | 0.0005358 | 0.00045169 | 7.4655e-05 | 0.00044708
| 24.9 | 24.85 | 0.00058316 | 0.00049063 | 0.00037438 | 23 | 22.95 | 2401725.0907 | 2408699.7755 | 2390586.6898 | 46.9 | 46.8 | 0.0004314 | 23 | 22.95 | 3175012.3108 |  | 0.00053895 | 0.0004543 | 7.1273e-05 | 0.00044864
| 25 | 24.95 | 0.00058444 | 0.00049395 | 0.00037614 | 23.1 | 23.05 | 2415224.2895 | 2422153.2373 | 2404158.4247 | 47.1 | 47 | 0.00044284 | 23.1 | 23.05 | 3185611.3243 |  | 0.000542 | 0.00045684 | 7.3163e-05 | 0.00045014
| 25.1 | 25.05 | 0.0005876 | 0.00049379 | 0.0003788 | 23.2 | 23.15 | 2428255.8446 | 2435138.1809 | 2417263.9273 | 47.3 | 47.2 | 0.00040179 | 23.2 | 23.15 | 3195688.1458 |  | 0.00054496 | 0.00045929 | 6.6382e-05 | 0.00045157
| 25.2 | 25.15 | 0.00058327 | 0.00049145 | 0.00037594 | 23.3 | 23.25 | 2440861.7076 | 2447696.7174 | 2429944.8944 | 47.5 | 47.4 | 0.00038693 | 23.3 | 23.25 | 3205262.2964 |  | 0.00054782 | 0.00046166 | 6.3927e-05 | 0.00045292
| 25.3 | 25.25 | 0.00058529 | 0.00049377 | 0.00037942 | 23.4 | 23.35 | 2453046.4614 | 2459833.4787 | 2442205.8303 | 47.7 | 47.6 | 0.00033785 | 23.4 | 23.35 | 3214364.8157 |  | 0.00055058 | 0.00046395 | 5.5817e-05 | 0.0004542
| 25.4 | 25.35 | 0.0005876 | 0.00049543 | 0.00037803 | 23.5 | 23.45 | 2464850.095 | 2471588.5979 | 2454086.4943 | 47.9 | 47.8 | 0.00026258 | 23.5 | 23.45 | 3222984.7924 |  | 0.00055326 | 0.00046617 | 4.3382e-05 | 0.00045542
| 25.5 | 25.45 | 0.00059105 | 0.00049873 | 0.00037878 | 23.6 | 23.55 | 2476237.7539 | 2482927.1556 | 2465552.1345 | 48.1 | 48 | 0.00024035 | 23.6 | 23.55 | 3231200.8601 |  | 0.00055585 | 0.0004683 | 3.9709e-05 | 0.00045658
| 25.6 | 25.55 | 0.00059415 | 0.00050114 | 0.00038325 | 23.7 | 23.65 | 2487177.4328 | 2493817.0929 | 2476570.8316 | 48.3 | 48.2 | 0.00019217 | 23.7 | 23.65 | 3238924.6872 |  | 0.00055833 | 0.00047036 | 3.1749e-05 | 0.00045768
| 25.7 | 25.65 | 0.0005941 | 0.00050265 | 0.00038308 | 23.8 | 23.75 | 2497710.0863 | 2504299.5063 | 2487183.3135 | 48.5 | 48.4 | 0.00014709 | 23.8 | 23.75 | 3246159.8236 |  | 0.00056072 | 0.00047234 | 2.4302e-05 | 0.0004587
| 25.8 | 25.75 | 0.0005916 | 0.00049638 | 0.00037947 | 23.9 | 23.85 | 2507836.9306 | 2514375.6434 | 2497390.745 | 48.7 | 48.6 | 9.954e-05 | 23.9 | 23.85 | 3252956.2377 |  | 0.00056302 | 0.00047424 | 1.6445e-05 | 0.00045966
| 25.9 | 25.85 | 0.00059025 | 0.00049924 | 0.00038057 | 24 | 23.95 | 2517611.0534 | 2524098.7567 | 2507245.9515 | 48.9 | 48.8 | 9.4562e-06 | 24 | 23.95 | 3259309.2486 |  | 0.00056525 | 0.00047607 | 1.5623e-06 | 0.00046056
| 26 | 25.95 | 0.00059323 | 0.00050112 | 0.00038232 | 24.1 | 24.05 | 2526952.9606 | 2533389.1712 | 2516669.727 | 49.1 | 49 | 0 | 24.1 | 24.05 | 3265284.5007 |  | 0.00056737 | 0.00047782 | 0 | 0.0004614
| 26.1 | 26.05 | 0.00058944 | 0.00049653 | 0.00038091 | 24.2 | 24.15 | 2535882.2886 | 2542266.601 | 2525681.584 | 49.3 | 49.2 | 0 | 24.2 | 24.15 | 3270790.2897 |  | 0.0005694 | 0.0004795 | 0 | 0.00046218
| 26.2 | 26.15 | 0.00059332 | 0.00050074 | 0.0003808 | 24.3 | 24.25 | 2544412.9384 | 2550745.0068 | 2534295.3275 | 49.5 | 49.4 | 0 | 24.3 | 24.25 | 3275854.8025 |  | 0.00057134 | 0.0004811 | 0 | 0.00046289
| 26.3 | 26.25 | 0.00059055 | 0.00049911 | 0.00038253 | 24.4 | 24.35 | 2552558.1543 | 2558837.6889 | 2542524.1124 |  |  |  | 24.4 | 24.35 | 3280500.3311 |  | 0.0005732 | 0.00048262 | 0 | 0.00046355
| 26.4 | 26.35 | 0.00059257 | 0.00049959 | 0.00038284 | 24.5 | 24.45 | 2560385.6009 | 2566612.4978 | 2550435.3061 |  |  |  | 24.5 | 24.45 | 3284734.3563 |  | 0.00057498 | 0.00048409 | 0 | 0.00046415
| 26.5 | 26.45 | 0.0005915 | 0.00049909 | 0.00037963 | 24.6 | 24.55 | 2567766.0943 | 2573939.9537 | 2557900.1967 |  |  |  | 24.6 | 24.55 | 3288636.8646 |  | 0.00057667 | 0.00048547 | 0 | 0.0004647
| 26.6 | 26.55 | 0.00059542 | 0.00050097 | 0.00038421 | 24.7 | 24.65 | 2574768.1694 | 2580888.7797 | 2564987.0191 |  |  |  | 24.7 | 24.65 | 3292094.3305 |  | 0.00057826 | 0.00048678 | 0 | 0.00046519
| 26.7 | 26.65 | 0.00059252 | 0.00050097 | 0.00038304 | 24.8 | 24.75 | 2581388.6208 | 2587455.781 | 2571692.5498 |  |  |  | 24.8 | 24.75 | 3295132.059 |  | 0.00057977 | 0.00048802 | 0 | 0.00046562
| 26.8 | 26.75 | 0.00059277 | 0.00050041 | 0.00038245 | 24.9 | 24.85 | 2587663.165 | 2593676.7756 | 2578052.3436 |  |  |  | 24.9 | 24.85 | 3297782.9515 |  | 0.00058121 | 0.00048919 | 0 | 0.00046599
| 26.9 | 26.85 | 0.00059671 | 0.00050342 | 0.00038464 | 25 | 24.95 | 2593622.9722 | 2599583.0198 | 2584097.4326 |  |  |  | 25 | 24.95 | 3300056.1801 |  | 0.00058257 | 0.00049031 | 0 | 0.00046631
| 27 | 26.95 | 0.00059393 | 0.00050202 | 0.00038313 | 25.1 | 25.05 | 2599147.5879 | 2605053.7972 | 2589707.7808 |  |  |  | 25.1 | 25.05 | 3302022.3239 |  | 0.00058384 | 0.00049134 | 0 | 0.00046659
| 27.1 | 27.05 | 0.00059532 | 0.00049965 | 0.00038261 | 25.2 | 25.15 | 2604003.3591 | 2609855.7299 | 2594649.2987 |  |  |  | 25.2 | 25.15 | 3303235.6591 |  | 0.00058495 | 0.00049224 | 0 | 0.00046676
| 27.2 | 27.15 | 0.00059813 | 0.00050562 | 0.00038561 | 25.3 | 25.25 | 2608178.6281 | 2613977.145 | 2598910.3528 |  |  |  | 25.3 | 25.25 | 3303743.7291 |  | 0.00058591 | 0.00049302 | 0 | 0.00046683
| 27.3 | 27.25 | 0.00059396 | 0.00050097 | 0.00038261 | 25.4 | 25.35 | 2612060.6319 | 2617805.3787 | 2602878.0171 |  |  |  | 25.4 | 25.35 | 3303899.4787 |  | 0.00058681 | 0.00049374 | 0 | 0.00046686
| 27.4 | 27.35 | 0.00059448 | 0.00050042 | 0.00038327 | 25.5 | 25.45 | 2615628.5728 | 2621319.596 | 2606531.5537 |  |  |  | 25.5 | 25.45 | 3303721.7456 |  | 0.00058763 | 0.00049441 | 0 | 0.00046683
| 27.5 | 27.45 | 0.00059027 | 0.00049725 | 0.00038157 | 25.6 | 25.55 | 2618810.8942 | 2624448.0952 | 2609799.6377 |  |  |  | 25.6 | 25.55 | 3303291.9615 |  | 0.00058837 | 0.000495 | 0 | 0.00046677
| 27.6 | 27.55 | 0.00059519 | 0.00050235 | 0.00038357 | 25.7 | 25.65 | 2621658.9738 | 2627242.377 | 2612733.4501 |  |  |  | 25.7 | 25.65 | 3302478.1871 |  | 0.00058903 | 0.00049552 | 0 | 0.00046666
| 27.7 | 27.65 | 0.00058966 | 0.00049827 | 0.0003805 | 25.8 | 25.75 | 2624202.1312 | 2629731.8322 | 2615362.1962 |  |  |  | 25.8 | 25.75 | 3301323.2966 |  | 0.00058962 | 0.00049599 | 0 | 0.00046649
| 27.8 | 27.75 | 0.00059586 | 0.00050326 | 0.00038644 | 25.9 | 25.85 | 2626483.6243 | 2631959.8161 | 2617728.9781 |  |  |  | 25.9 | 25.85 | 3299848.8128 |  | 0.00059015 | 0.00049641 | 0 | 0.00046628
| 27.9 | 27.85 | 0.00058884 | 0.0004974 | 0.00038066 | 26 | 25.95 | 2628429.2428 | 2633851.9695 | 2619759.8235 |  |  |  | 26 | 25.95 | 3298071.5439 |  | 0.00059061 | 0.00049677 | 0 | 0.00046603
| 28 | 27.95 | 0.00058761 | 0.00049554 | 0.00037946 | 26.1 | 26.05 | 2630034.9988 | 2635404.3051 | 2621450.7434 |  |  |  | 26.1 | 26.05 | 3296075.0712 |  | 0.00059099 | 0.00049706 | 0 | 0.00046575
| 28.1 | 28.05 | 0.0005878 | 0.00049502 | 0.00037779 | 26.2 | 26.15 | 2631317.5993 | 2636633.572 | 2622818.377 |  |  |  | 26.2 | 26.15 | 3293722.0309 |  | 0.0005913 | 0.0004973 | 0 | 0.00046542
| 28.2 | 28.15 | 0.00059075 | 0.00049914 | 0.0003836 | 26.3 | 26.25 | 2632322.5916 | 2637585.416 | 2623908.1143 |  |  |  | 26.3 | 26.25 | 3291052.8978 |  | 0.00059155 | 0.00049747 | 0 | 0.00046504
| 28.3 | 28.25 | 0.00058473 | 0.00049303 | 0.00037884 | 26.4 | 26.35 | 2633108.4308 | 2638318.4118 | 2624778.2186 |  |  |  | 26.4 | 26.35 | 3288081.6766 |  | 0.00059174 | 0.00049761 | 0 | 0.00046462
| 28.4 | 28.35 | 0.00058819 | 0.00049683 | 0.00038 | 26.5 | 26.45 | 2633531.3545 | 2638688.5159 | 2625285.3772 |  |  |  | 26.5 | 26.45 | 3284822.6956 |  | 0.00059186 | 0.00049768 | 0 | 0.00046416
| 28.5 | 28.45 | 0.00058195 | 0.0004903 | 0.00037583 | 26.6 | 26.55 | 2633651.8787 | 2638756.3704 | 2625489.904 |  |  |  | 26.6 | 26.55 | 3281382.769 |  | 0.0005919 | 0.0004977 | 0 | 0.00046367
| 28.6 | 28.55 | 0.00058561 | 0.00049382 | 0.00037893 | 26.7 | 26.65 | 2633487.6055 | 2638539.6162 | 2625409.339 |  |  |  | 26.7 | 26.65 | 3277603.4328 |  | 0.00059188 | 0.00049765 | 0 | 0.00046314
| 28.7 | 28.65 | 0.00058231 | 0.00049256 | 0.00037736 | 26.8 | 26.75 | 2633061.6237 | 2638061.3896 | 2625066.695 |  |  |  | 26.8 | 26.75 | 3273522.7964 |  | 0.00059181 | 0.00049756 | 0 | 0.00046256
| 28.8 | 28.75 | 0.00058389 | 0.00049291 | 0.00037853 | 26.9 | 26.85 | 2632435.3475 | 2637383.2224 | 2624523.1977 |  |  |  | 26.9 | 26.85 | 3269169.3653 |  | 0.00059169 | 0.00049744 | 0 | 0.00046195
| 28.9 | 28.85 | 0.00058031 | 0.00049054 | 0.00037505 | 27 | 26.95 | 2631447.5337 | 2636343.57 | 2623618.0859 |  |  |  | 27 | 26.95 | 3264574.3966 |  | 0.00059148 | 0.00049724 | 0 | 0.0004613
| 29 | 28.95 | 0.00057996 | 0.00048822 | 0.00037359 | 27.1 | 27.05 | 2630176.3362 | 2635020.7376 | 2622429.2711 |  |  |  | 27.1 | 27.05 | 3259810.9213 |  | 0.00059121 | 0.00049699 | 0 | 0.00046063
| 29.1 | 29.05 | 0.00057551 | 0.00048549 | 0.00037185 | 27.2 | 27.15 | 2628653.072 | 2633446.1027 | 2620987.9742 |  |  |  | 27.2 | 27.15 | 3254738.2286 |  | 0.00059089 | 0.00049669 | 0 | 0.00045991
| 29.2 | 29.15 | 0.00057691 | 0.0004859 | 0.00037023 | 27.3 | 27.25 | 2626918.8361 | 2631660.835 | 2619335.1694 |  |  |  | 27.3 | 27.25 | 3249389.7066 |  | 0.00059052 | 0.00049636 | 0 | 0.00045915
| 29.3 | 29.25 | 0.00057284 | 0.00048256 | 0.00037172 | 27.4 | 27.35 | 2624946.8805 | 2629638.1379 | 2617444.1874 |  |  |  | 27.4 | 27.35 | 3243778.4573 |  | 0.00059009 | 0.00049598 | 0 | 0.00045836
| 29.4 | 29.35 | 0.00057789 | 0.00048613 | 0.00037243 | 27.5 | 27.45 | 2622657.3763 | 2627298.0412 | 2615235.4246 |  |  |  | 27.5 | 27.45 | 3237957.1735 |  | 0.00058959 | 0.00049553 | 0 | 0.00045754
| 29.5 | 29.45 | 0.00057229 | 0.00048355 | 0.00036789 | 27.6 | 27.55 | 2619638.3628 | 2624229.1069 | 2612296.0884 |  |  |  | 27.6 | 27.55 | 3231534.4671 |  | 0.00058893 | 0.00049496 | 0 | 0.00045663
| 29.6 | 29.55 | 0.00057517 | 0.00048564 | 0.0003729 | 27.7 | 27.65 | 2616147.2987 | 2620688.6857 | 2608883.8084 |  |  |  | 27.7 | 27.65 | 3224563.467 |  | 0.00058816 | 0.00049429 | 0 | 0.00045565
| 29.7 | 29.65 | 0.00056878 | 0.00048155 | 0.00036991 | 27.8 | 27.75 | 2612502.1046 | 2616994.5468 | 2605316.7445 |  |  |  | 27.8 | 27.75 | 3217360.5972 |  | 0.00058736 | 0.00049359 | 0 | 0.00045463
| 29.8 | 29.75 | 0.00057117 | 0.00048178 | 0.00037037 | 27.9 | 27.85 | 2608615.4611 | 2613059.2177 | 2601507.8216 |  |  |  | 27.9 | 27.85 | 3209949.8065 |  | 0.0005865 | 0.00049285 | 0 | 0.00045358
| 29.9 | 29.85 | 0.00056819 | 0.00048094 | 0.00036722 | 28 | 27.95 | 2604456.6011 | 2608851.8797 | 2597426.3554 |  |  |  | 28 | 27.95 | 3202350.6118 |  | 0.00058558 | 0.00049206 | 0 | 0.00045251
| 30 | 29.95 | 0.00057214 | 0.00048449 | 0.00036959 | 28.1 | 28.05 | 2600078.6571 | 2604425.7556 | 2593125.3337 |  |  |  | 28.1 | 28.05 | 3194657.2977 |  | 0.00058461 | 0.00049122 | 0 | 0.00045142
| 30.1 | 30.05 | 0.00056568 | 0.0004771 | 0.0003655 | 28.2 | 28.15 | 2595509.5696 | 2599808.831 | 2588632.6247 |  |  |  | 28.2 | 28.15 | 3186706.3376 |  | 0.00058359 | 0.00049035 | 0 | 0.0004503
| 30.2 | 30.15 | 0.00056866 | 0.00048024 | 0.00036609 | 28.3 | 28.25 | 2590822.8744 | 2595074.7602 | 2584021.5749 |  |  |  | 28.3 | 28.25 | 3178535.2508 |  | 0.00058255 | 0.00048946 | 0 | 0.00044914
| 30.3 | 30.25 | 0.00056264 | 0.00047475 | 0.00036464 | 28.4 | 28.35 | 2585861.096 | 2590065.8056 | 2579135.1274 |  |  |  | 28.4 | 28.35 | 3170164.2517 |  | 0.00058145 | 0.00048851 | 0 | 0.00044796
| 30.4 | 30.35 | 0.00056049 | 0.00047363 | 0.00036219 | 28.5 | 28.45 | 2580662.0166 | 2584819.812 | 2574010.9645 |  |  |  | 28.5 | 28.45 | 3161639.2704 |  | 0.0005803 | 0.00048752 | 0 | 0.00044675
| 30.5 | 30.45 | 0.00055886 | 0.00047006 | 0.00036024 | 28.6 | 28.55 | 2575277.6842 | 2579388.9099 | 2568701.0019 |  |  |  | 28.6 | 28.55 | 3153043.3332 |  | 0.0005791 | 0.0004865 | 0 | 0.00044554
| 30.6 | 30.55 | 0.00056033 | 0.00047405 | 0.00036389 | 28.7 | 28.65 | 2569733.4674 | 2573798.5054 | 2563230.5481 |  |  |  | 28.7 | 28.65 | 3144192.3073 |  | 0.00057787 | 0.00048544 | 0 | 0.00044429
| 30.7 | 30.65 | 0.00055983 | 0.00047216 | 0.00036189 | 28.8 | 28.75 | 2564075.5118 | 2568094.8129 | 2557645.639 |  |  |  | 28.8 | 28.75 | 3135138.0601 |  | 0.00057661 | 0.00048437 | 0 | 0.00044301
| 30.8 | 30.75 | 0.00055671 | 0.00046846 | 0.00035992 | 28.9 | 28.85 | 2558129.5764 | 2562103.3149 | 2551772.4753 |  |  |  | 28.9 | 28.85 | 3125905.8632 |  | 0.00057528 | 0.00048324 | 0 | 0.0004417
| 30.9 | 30.85 | 0.00055162 | 0.00046649 | 0.00035761 | 29 | 28.95 | 2551968.7408 | 2555897.2053 | 2545683.954 |  |  |  | 29 | 28.95 | 3116531.2834 |  | 0.00057391 | 0.00048207 | 0 | 0.00044038
| 31 | 30.95 | 0.00055171 | 0.00046704 | 0.00035718 | 29.1 | 29.05 | 2545663.2265 | 2549546.8112 | 2539450.1272 |  |  |  | 29.1 | 29.05 | 3107107.0795 |  | 0.00057251 | 0.00048087 | 0 | 0.00043905
| 31.1 | 31.05 | 0.00055241 | 0.00046266 | 0.00035563 | 29.2 | 29.15 | 2539225.8145 | 2543064.9288 | 2533083.7517 |  |  |  | 29.2 | 29.15 | 3097442.0039 |  | 0.00057107 | 0.00047965 | 0 | 0.00043768
| 31.2 | 31.15 | 0.00054749 | 0.00046194 | 0.00035447 | 29.3 | 29.25 | 2532637.3646 | 2536432.3843 | 2526565.7406 |  |  |  | 29.3 | 29.25 | 3087585.0506 |  | 0.0005696 | 0.0004784 | 0 | 0.00043629
| 31.3 | 31.25 | 0.00054596 | 0.00046196 | 0.00035257 | 29.4 | 29.35 | 2525780.2095 | 2529531.332 | 2519778.7126 |  |  |  | 29.4 | 29.35 | 3077581.6682 |  | 0.00056807 | 0.00047709 | 0 | 0.00043488
| 31.4 | 31.35 | 0.00054593 | 0.00046107 | 0.00035216 | 29.5 | 29.45 | 2518761.1708 | 2522468.7513 | 2512829.2362 |  |  |  | 29.5 | 29.45 | 3067451.4638 |  | 0.0005665 | 0.00047576 | 0 | 0.00043345
| 31.5 | 31.45 | 0.00054024 | 0.00045443 | 0.00034872 | 29.6 | 29.55 | 2511592.5887 | 2515256.9971 | 2505729.6283 |  |  |  | 29.6 | 29.55 | 3057293.943 |  | 0.0005649 | 0.0004744 | 0 | 0.00043201
| 31.6 | 31.55 | 0.00053783 | 0.00045491 | 0.00034764 | 29.7 | 29.65 | 2504320.9501 | 2507942.6196 | 2498526.2744 |  |  |  | 29.7 | 29.65 | 3046900.2402 |  | 0.00056328 | 0.00047302 | 0 | 0.00043054
| 31.7 | 31.65 | 0.00053734 | 0.00045137 | 0.00034598 | 29.8 | 29.75 | 2496870.902 | 2500450.1521 | 2491144.003 |  |  |  | 29.8 | 29.75 | 3036329.9666 |  | 0.00056162 | 0.00047161 | 0 | 0.00042905
| 31.8 | 31.75 | 0.00053706 | 0.00045287 | 0.00034622 | 29.9 | 29.85 | 2489198.5885 | 2492735.6737 | 2483539.0623 |  |  |  | 29.9 | 29.85 | 3025633.0634 |  | 0.0005599 | 0.00047015 | 0 | 0.00042754
| 31.9 | 31.85 | 0.00053519 | 0.00045429 | 0.00034802 | 30 | 29.95 | 2481374.8918 | 2484870.1655 | 2475782.1763 |  |  |  | 30 | 29.95 | 3014832.3914 |  | 0.00055815 | 0.00046867 | 0 | 0.00042601
| 32 | 31.95 | 0.0005315 | 0.00044829 | 0.00034332 | 30.1 | 30.05 | 2473421.8036 | 2476875.6463 | 2467895.2933 |  |  |  | 30.1 | 30.05 | 3004016.9901 |  | 0.00055637 | 0.00046716 | 0 | 0.00042448
| 32.1 | 32.05 | 0.00052874 | 0.0004468 | 0.00034052 | 30.2 | 30.15 | 2465116.0678 | 2468529.3069 | 2459654.4441 |  |  |  | 30.2 | 30.15 | 2992697.0811 |  | 0.00055452 | 0.00046559 | 0 | 0.00042288
| 32.2 | 32.15 | 0.00053217 | 0.00044791 | 0.00034211 | 30.3 | 30.25 | 2456330.5129 | 2459703.7739 | 2450932.7798 |  |  |  | 30.3 | 30.25 | 2980940.3234 |  | 0.00055255 | 0.00046392 | 0 | 0.00042122
| 32.3 | 32.25 | 0.00052221 | 0.00044067 | 0.00033952 | 30.4 | 30.35 | 2447367.737 | 2450701.2949 | 2442033.4573 |  |  |  | 30.4 | 30.35 | 2969108.7933 |  | 0.00055054 | 0.00046223 | 0 | 0.00041955
| 32.4 | 32.35 | 0.00052274 | 0.00044064 | 0.00033671 | 30.5 | 30.45 | 2438300.0622 | 2441594.2876 | 2433028.6457 |  |  |  | 30.5 | 30.45 | 2957200.2065 |  | 0.00054851 | 0.00046051 | 0 | 0.00041787
| 32.5 | 32.45 | 0.000519 | 0.00043897 | 0.00033595 | 30.6 | 30.55 | 2429129.0009 | 2432384.2617 | 2423919.8613 |  |  |  | 30.6 | 30.55 | 2945289.3868 |  | 0.00054646 | 0.00045877 | 0 | 0.00041618
| 32.6 | 32.55 | 0.00051697 | 0.00043622 | 0.00033398 | 30.7 | 30.65 | 2419923.1547 | 2423139.9057 | 2414775.567 |  |  |  | 30.7 | 30.65 | 2933186.2926 |  | 0.0005444 | 0.00045703 | 0 | 0.00041447
| 32.7 | 32.65 | 0.00051538 | 0.00043396 | 0.0003317 | 30.8 | 30.75 | 2410504.375 | 2413682.8293 | 2405418.0009 |  |  |  | 30.8 | 30.75 | 2920951.7348 |  | 0.00054229 | 0.00045524 | 0 | 0.00041274
| 32.8 | 32.75 | 0.00051614 | 0.00043493 | 0.00033193 | 30.9 | 30.85 | 2400945.2829 | 2404085.7477 | 2395919.633 |  |  |  | 30.9 | 30.85 | 2908649.347 |  | 0.00054015 | 0.00045343 | 0 | 0.00041101
| 32.9 | 32.85 | 0.00050798 | 0.00042906 | 0.00032881 | 31 | 30.95 | 2391305.0786 | 2394407.9348 | 2386339.546 |  |  |  | 31 | 30.95 | 2896264.3441 |  | 0.00053799 | 0.00045161 | 0 | 0.00040926
| 33 | 32.95 | 0.00050979 | 0.00042999 | 0.00032932 | 31.1 | 31.05 | 2381589.809 | 2384655.4402 | 2376683.782 |  |  |  | 31.1 | 31.05 | 2883889.8637 |  | 0.00053581 | 0.00044977 | 0 | 0.00040751
| 33.1 | 33.05 | 0.0005066 | 0.00042632 | 0.00032532 | 31.2 | 31.15 | 2371818.7499 | 2374847.5593 | 2366971.5854 |  |  |  | 31.2 | 31.15 | 2871337.105 |  | 0.00053362 | 0.00044792 | 0 | 0.00040573
| 33.2 | 33.15 | 0.00050251 | 0.0004243 | 0.0003266 | 31.3 | 31.25 | 2361824.4981 | 2364816.6713 | 2357035.9016 |  |  |  | 31.3 | 31.25 | 2858665.2438 |  | 0.00053138 | 0.00044603 | 0 | 0.00040394
| 33.3 | 33.25 | 0.00050277 | 0.00042291 | 0.000324 | 31.4 | 31.35 | 2351733.1238 | 2354689.0044 | 2347002.5477 |  |  |  | 31.4 | 31.35 | 2845936.0886 |  | 0.00052912 | 0.00044412 | 0 | 0.00040214
| 33.4 | 33.35 | 0.00049942 | 0.0004204 | 0.00032182 | 31.5 | 31.45 | 2341561.5151 | 2344481.4634 | 2336888.3847 |  |  |  | 31.5 | 31.45 | 2833150.7844 |  | 0.00052684 | 0.00044219 | 0 | 0.00040034
| 33.5 | 33.45 | 0.00049682 | 0.00041955 | 0.0003197 | 31.6 | 31.55 | 2331348.5439 | 2334232.9638 | 2326732.2151 |  |  |  | 31.6 | 31.55 | 2820375.4094 |  | 0.00052455 | 0.00044026 | 0 | 0.00039853
| 33.6 | 33.55 | 0.00049405 | 0.00041716 | 0.00032134 | 31.7 | 31.65 | 2321029.503 | 2323878.7132 | 2316469.4677 |  |  |  | 31.7 | 31.65 | 2807442.7755 |  | 0.00052224 | 0.00043831 | 0 | 0.0003967
| 33.7 | 33.65 | 0.00048949 | 0.0004133 | 0.00031651 | 31.8 | 31.75 | 2310525.2739 | 2313339.493 | 2306021.1843 |  |  |  | 31.8 | 31.75 | 2794402.412 |  | 0.00051988 | 0.00043632 | 0 | 0.00039486
| 33.8 | 33.75 | 0.00049072 | 0.00041336 | 0.00031662 | 31.9 | 31.85 | 2299948.3667 | 2302727.9479 | 2295499.6598 |  |  |  | 31.9 | 31.85 | 2781315.2046 |  | 0.00051751 | 0.00043432 | 0 | 0.00039301
| 33.9 | 33.85 | 0.00048446 | 0.00040853 | 0.00031214 | 32 | 31.95 | 2289296.2493 | 2292041.5379 | 2284902.3739 |  |  |  | 32 | 31.95 | 2768176.7658 |  | 0.00051512 | 0.0004323 | 0 | 0.00039116
| 34 | 33.95 | 0.00048633 | 0.00041077 | 0.00031283 | 32.1 | 32.05 | 2278635.7882 | 2281347.2049 | 2274296.073 |  |  |  | 32.1 | 32.05 | 2755065.7924 |  | 0.00051273 | 0.00043028 | 0 | 0.0003893
| 34.1 | 34.05 | 0.00047999 | 0.00040364 | 0.00030875 | 32.2 | 32.15 | 2267821.6714 | 2270499.4591 | 2263535.7298 |  |  |  | 32.2 | 32.15 | 2741798.9639 |  | 0.0005103 | 0.00042824 | 0 | 0.00038743
| 34.2 | 34.15 | 0.00047725 | 0.00040299 | 0.00030808 | 32.3 | 32.25 | 2256875.4935 | 2259519.9179 | 2252642.9015 |  |  |  | 32.3 | 32.25 | 2728431.7446 |  | 0.00050785 | 0.00042617 | 0 | 0.00038554
| 34.3 | 34.25 | 0.0004753 | 0.00040156 | 0.00030761 | 32.4 | 32.35 | 2245870.1339 | 2248481.5432 | 2241690.3364 |  |  |  | 32.4 | 32.35 | 2715032.397 |  | 0.00050538 | 0.00042409 | 0 | 0.00038365
| 34.4 | 34.35 | 0.00047311 | 0.00039706 | 0.00030309 | 32.5 | 32.45 | 2234788.5402 | 2237367.258 | 2230661.021 |  |  |  | 32.5 | 32.45 | 2701591.5194 |  | 0.00050289 | 0.00042199 | 0 | 0.00038175
| 34.5 | 34.45 | 0.0004663 | 0.00039638 | 0.00030184 | 32.6 | 32.55 | 2223702.8674 | 2226249.2959 | 2219626.9843 |  |  |  | 32.6 | 32.55 | 2688176.529 |  | 0.0005004 | 0.00041989 | 0 | 0.00037985
| 34.6 | 34.55 | 0.00046592 | 0.0003921 | 0.0003007 | 32.7 | 32.65 | 2212444.5472 | 2214958.891 | 2208419.9743 |  |  |  | 32.7 | 32.65 | 2674609.5073 |  | 0.00049788 | 0.00041776 | 0 | 0.00037793
| 34.7 | 34.65 | 0.00045881 | 0.00038555 | 0.00029456 | 32.8 | 32.75 | 2201091.5572 | 2203574.1076 | 2197117.83 |  |  |  | 32.8 | 32.75 | 2660960.2777 |  | 0.00049533 | 0.00041562 | 0 | 0.00037601
| 34.8 | 34.75 | 0.00046054 | 0.0003883 | 0.00029787 | 32.9 | 32.85 | 2189685.1423 | 2192136.2332 | 2185761.7279 |  |  |  | 32.9 | 32.85 | 2647281.2094 |  | 0.00049277 | 0.00041346 | 0 | 0.00037407
| 34.9 | 34.85 | 0.00045609 | 0.00038469 | 0.00029352 | 33 | 32.95 | 2178237.4099 | 2180657.3842 | 2174363.7609 |  |  |  | 33 | 32.95 | 2633572.158 |  | 0.0004902 | 0.00041129 | 0 | 0.00037214
| 35 | 34.95 | 0.00045629 | 0.0003867 | 0.00029497 | 33.1 | 33.05 | 2166759.9203 | 2169149.1289 | 2162935.4763 |  |  |  | 33.1 | 33.05 | 2619892.9729 |  | 0.00048762 | 0.00040912 | 0 | 0.0003702
| 35.1 | 35.05 | 0.000455 | 0.00038031 | 0.00029037 | 33.2 | 33.15 | 2155115.0845 | 2157473.7231 | 2151339.5341 |  |  |  | 33.2 | 33.15 | 2606063.4043 |  | 0.00048501 | 0.00040692 | 0 | 0.00036825
| 35.2 | 35.15 | 0.00044906 | 0.00037851 | 0.00028818 | 33.3 | 33.25 | 2143406.7183 | 2145735.0938 | 2139679.5716 |  |  |  | 33.3 | 33.25 | 2592159.7091 |  | 0.00048238 | 0.00040471 | 0 | 0.00036628
| 35.3 | 35.25 | 0.00044305 | 0.0003739 | 0.00028598 | 33.4 | 33.35 | 2131650.0879 | 2133948.5197 | 2127970.8351 |  |  |  | 33.4 | 33.35 | 2578248.9425 |  | 0.00047974 | 0.00040248 | 0 | 0.00036432
| 35.4 | 35.35 | 0.00044465 | 0.00037287 | 0.00028483 | 33.5 | 33.45 | 2119884.6112 | 2122153.4568 | 2116252.6815 |  |  |  | 33.5 | 33.45 | 2564316.483 |  | 0.0004771 | 0.00040026 | 0 | 0.00036235
| 35.5 | 35.45 | 0.00043813 | 0.00037223 | 0.00028141 | 33.6 | 33.55 | 2108052.8791 | 2110292.4296 | 2104467.8076 |  |  |  | 33.6 | 33.55 | 2550413.0503 |  | 0.00047444 | 0.00039802 | 0 | 0.00036039
| 35.6 | 35.55 | 0.00043505 | 0.00036675 | 0.00028002 | 33.7 | 33.65 | 2096071.5642 | 2098282.0194 | 2092533.0324 |  |  |  | 33.7 | 33.65 | 2536365.5228 |  | 0.00047175 | 0.00039576 | 0 | 0.0003584
| 35.7 | 35.65 | 0.00043399 | 0.00036469 | 0.00027888 | 33.8 | 33.75 | 2084048.5563 | 2086230.2273 | 2080556.0677 |  |  |  | 33.8 | 33.75 | 2522265.8597 |  | 0.00046905 | 0.00039348 | 0 | 0.00035641
| 35.8 | 35.75 | 0.00043029 | 0.00036275 | 0.00027732 | 33.9 | 33.85 | 2071971.7599 | 2074124.941 | 2068524.8446 |  |  |  | 33.9 | 33.85 | 2508156.3238 |  | 0.00046634 | 0.0003912 | 0 | 0.00035441
| 35.9 | 35.85 | 0.00042467 | 0.00035993 | 0.00027444 | 34 | 33.95 | 2059910.9663 | 2062036.0201 | 2056509.0452 |  |  |  | 34 | 33.95 | 2494026.3298 |  | 0.00046363 | 0.00038892 | 0 | 0.00035242
| 36 | 35.95 | 0.0004233 | 0.00035595 | 0.00027331 | 34.1 | 34.05 | 2047728.8095 | 2049825.9523 | 2044371.5371 |  |  |  | 34.1 | 34.05 | 2479922.4025 |  | 0.00046089 | 0.00038662 | 0 | 0.00035042
| 36.1 | 36.05 | 0.00041721 | 0.00035192 | 0.00026771 | 34.2 | 34.15 | 2035447.4235 | 2037516.8924 | 2032134.4219 |  |  |  | 34.2 | 34.15 | 2465689.1363 |  | 0.00045813 | 0.0003843 | 0 | 0.00034841
| 36.2 | 36.15 | 0.00041939 | 0.00035304 | 0.00027013 | 34.3 | 34.25 | 2023133.2736 | 2025175.3698 | 2019864.0615 |  |  |  | 34.3 | 34.25 | 2451408.8571 |  | 0.00045537 | 0.00038197 | 0 | 0.0003464
| 36.3 | 36.25 | 0.00041172 | 0.0003457 | 0.00026419 | 34.4 | 34.35 | 2010784.1112 | 2012799.1298 | 2007558.2173 |  |  |  | 34.4 | 34.35 | 2437120.1118 |  | 0.00045259 | 0.00037963 | 0 | 0.00034438
| 36.4 | 36.35 | 0.00041069 | 0.00034542 | 0.00026322 | 34.5 | 34.45 | 1998452.1905 | 2000440.4748 | 1995269.0668 |  |  |  | 34.5 | 34.45 | 2422821.1401 |  | 0.00044982 | 0.0003773 | 0 | 0.00034236
| 36.5 | 36.45 | 0.00040306 | 0.00033836 | 0.00025868 | 34.6 | 34.55 | 1985993.781 | 1987955.527 | 1982853.1143 |  |  |  | 34.6 | 34.55 | 2408552.5122 |  | 0.00044702 | 0.00037495 | 0 | 0.00034034
| 36.6 | 36.55 | 0.00040282 | 0.00033901 | 0.00025857 | 34.7 | 34.65 | 1973461.0456 | 1975396.4988 | 1970362.4439 |  |  |  | 34.7 | 34.65 | 2394164.2228 |  | 0.00044421 | 0.00037258 | 0 | 0.00033831
| 36.7 | 36.65 | 0.00039775 | 0.00033567 | 0.00025685 | 34.8 | 34.75 | 1960906.4042 | 1962815.8584 | 1957849.3986 |  |  |  | 34.8 | 34.75 | 2379730.5759 |  | 0.00044139 | 0.00037021 | 0 | 0.00033627
| 36.8 | 36.75 | 0.00039761 | 0.00033397 | 0.00025483 | 34.9 | 34.85 | 1948327.4032 | 1950211.1455 | 1945311.5347 |  |  |  | 34.9 | 34.85 | 2365304.7937 |  | 0.00043856 | 0.00036783 | 0 | 0.00033423
| 36.9 | 36.85 | 0.00039131 | 0.00032994 | 0.00025084 | 35 | 34.95 | 1935727.9172 | 1937586.2349 | 1932752.727 |  |  |  | 35 | 34.95 | 2350874.247 |  | 0.00043573 | 0.00036545 | 0 | 0.00033219
| 37 | 36.95 | 0.00039095 | 0.00032873 | 0.00024886 | 35.1 | 35.05 | 1923010.2579 | 1924843.3411 | 1920075.4426 |  |  |  | 35.1 | 35.05 | 2336479.1065 |  | 0.00043287 | 0.00036304 | 0 | 0.00033016
| 37.1 | 37.05 | 0.00038375 | 0.00032257 | 0.00024563 | 35.2 | 35.15 | 1910042.7769 | 1911851.2417 | 1907147.3522 |  |  |  | 35.2 | 35.15 | 2321772.3384 |  | 0.00042996 | 0.00036059 | 0 | 0.00032808
| 37.2 | 37.15 | 0.00038663 | 0.00032383 | 0.00024693 | 35.3 | 35.25 | 1896850.8606 | 1898635.323 | 1893993.8424 |  |  |  | 35.3 | 35.25 | 2306838.8358 |  | 0.00042699 | 0.0003581 | 0 | 0.00032597
| 37.3 | 37.25 | 0.00037579 | 0.00031608 | 0.0002405 | 35.4 | 35.35 | 1883674.5673 | 1885435.3073 | 1880855.5077 |  |  |  | 35.4 | 35.35 | 2291911.9329 |  | 0.00042403 | 0.00035561 | 0 | 0.00032386
| 37.4 | 37.35 | 0.0003778 | 0.0003179 | 0.00024248 | 35.5 | 35.45 | 1870470.5319 | 1872207.7858 | 1867689.0535 |  |  |  | 35.5 | 35.45 | 2277005.5161 |  | 0.00042106 | 0.00035312 | 0 | 0.00032175
| 37.5 | 37.45 | 0.00037025 | 0.00031157 | 0.00023737 | 35.6 | 35.55 | 1857158.5745 | 1858872.5015 | 1854414.4232 |  |  |  | 35.6 | 35.55 | 2262127.0513 |  | 0.00041807 | 0.0003506 | 0 | 0.00031965
| 37.6 | 37.55 | 0.00036932 | 0.00030952 | 0.00023601 | 35.7 | 35.65 | 1843833.7104 | 1845524.5554 | 1841126.4948 |  |  |  | 35.7 | 35.65 | 2247149.8209 |  | 0.00041507 | 0.00034808 | 0 | 0.00031753
| 37.7 | 37.65 | 0.00036147 | 0.00030463 | 0.00023251 | 35.8 | 35.75 | 1830489.102 | 1832157.1007 | 1827818.4457 |  |  |  | 35.8 | 35.75 | 2232157.2507 |  | 0.00041207 | 0.00034556 | 0 | 0.00031541
| 37.8 | 37.75 | 0.00036228 | 0.000304 | 0.00023122 | 35.9 | 35.85 | 1817171.3774 | 1818816.8045 | 1814536.8411 |  |  |  | 35.9 | 35.85 | 2217186.1741 |  | 0.00040908 | 0.00034305 | 0 | 0.0003133
| 37.9 | 37.85 | 0.00035936 | 0.00030204 | 0.00022807 | 36 | 35.95 | 1803798.5379 | 1805421.5907 | 1801199.8062 |  |  |  | 36 | 35.95 | 2202240.7075 |  | 0.00040607 | 0.00034052 | 0 | 0.00031119
| 38 | 37.95 | 0.00035511 | 0.00029667 | 0.00022754 | 36.1 | 36.05 | 1790341.5202 | 1791942.368 | 1787778.3232 |  |  |  | 36.1 | 36.05 | 2187322.4729 |  | 0.00040305 | 0.00033798 | 0 | 0.00030908
| 38.1 | 38.05 | 0.00034802 | 0.00029401 | 0.00022247 | 36.2 | 36.15 | 1776882.8703 | 1778461.7537 | 1774354.8231 |  |  |  | 36.2 | 36.15 | 2172325.9228 |  | 0.00040002 | 0.00033544 | 0 | 0.00030696
| 38.2 | 38.15 | 0.00034796 | 0.00029175 | 0.00022195 | 36.3 | 36.25 | 1763401.7917 | 1764958.9301 | 1760908.5439 |  |  |  | 36.3 | 36.25 | 2157313.2466 |  | 0.00039699 | 0.00033289 | 0 | 0.00030484
| 38.3 | 38.25 | 0.00034289 | 0.00028723 | 0.00021813 | 36.4 | 36.35 | 1749964.8593 | 1751500.5277 | 1747505.9712 |  |  |  | 36.4 | 36.35 | 2142318.4458 |  | 0.00039397 | 0.00033035 | 0 | 0.00030272
| 38.4 | 38.35 | 0.00033855 | 0.000284 | 0.00021547 | 36.5 | 36.45 | 1736442.5347 | 1737956.8913 | 1734017.7535 |  |  |  | 36.5 | 36.45 | 2127342.9626 |  | 0.00039093 | 0.0003278 | 0 | 0.0003006
| 38.5 | 38.45 | 0.00033346 | 0.00027917 | 0.00021179 | 36.6 | 36.55 | 1722861.7038 | 1724354.9287 | 1720470.7419 |  |  |  | 36.6 | 36.55 | 2112390.3229 |  | 0.00038787 | 0.00032523 | 0 | 0.00029849
| 38.6 | 38.55 | 0.00033253 | 0.00027852 | 0.00021184 | 36.7 | 36.65 | 1709279.4681 | 1710751.7886 | 1706921.9622 |  |  |  | 36.7 | 36.65 | 2097365.3118 |  | 0.00038482 | 0.00032266 | 0 | 0.00029637
| 38.7 | 38.65 | 0.00032825 | 0.00027423 | 0.00020809 | 36.8 | 36.75 | 1695703.0761 | 1697154.7226 | 1693378.6577 |  |  |  | 36.8 | 36.75 | 2082329.1582 |  | 0.00038176 | 0.0003201 | 0 | 0.00029424
| 38.8 | 38.75 | 0.00032593 | 0.00027358 | 0.00020568 | 36.9 | 36.85 | 1682133.8088 | 1683565.0099 | 1679842.1122 |  |  |  | 36.9 | 36.85 | 2067309.5964 |  | 0.00037871 | 0.00031754 | 0 | 0.00029212
| 38.9 | 38.85 | 0.00032083 | 0.00026842 | 0.00020457 | 37 | 36.95 | 1668475.6703 | 1669886.5708 | 1666216.4647 |  |  |  | 37 | 36.95 | 2052324.6165 |  | 0.00037564 | 0.00031496 | 0 | 0.00029
| 39 | 38.95 | 0.00031852 | 0.00026638 | 0.00020179 | 37.1 | 37.05 | 1654782.9297 | 1656173.7183 | 1652555.9134 |  |  |  | 37.1 | 37.05 | 2037360.7041 |  | 0.00037256 | 0.00031237 | 0 | 0.00028789
| 39.1 | 39.05 | 0.0003114 | 0.00026326 | 0.00019834 | 37.2 | 37.15 | 1641089.6827 | 1642460.5745 | 1638894.5117 |  |  |  | 37.2 | 37.15 | 2022335.0251 |  | 0.00036948 | 0.00030978 | 0 | 0.00028577
| 39.2 | 39.15 | 0.00030911 | 0.0002606 | 0.00019765 | 37.3 | 37.25 | 1627409.0492 | 1628760.2675 | 1625245.3667 |  |  |  | 37.3 | 37.25 | 2007290.7843 |  | 0.0003664 | 0.0003072 | 0 | 0.00028364
| 39.3 | 39.25 | 0.00030263 | 0.00025501 | 0.00019243 | 37.4 | 37.35 | 1613711.8911 | 1615043.6323 | 1611579.3831 |  |  |  | 37.4 | 37.35 | 1992267.935 |  | 0.00036332 | 0.00030461 | 0 | 0.00028152
| 39.4 | 39.35 | 0.0003013 | 0.00025182 | 0.00019139 | 37.5 | 37.45 | 1599915.7089 | 1601228.0993 | 1597814.1736 |  |  |  | 37.5 | 37.45 | 1977271.3454 |  | 0.00036022 | 0.00030201 | 0 | 0.0002794
| 39.5 | 39.45 | 0.00029908 | 0.0002518 | 0.0001906 | 37.6 | 37.55 | 1586121.5418 | 1587414.789 | 1584050.6473 |  |  |  | 37.6 | 37.55 | 1962294.4647 |  | 0.00035712 | 0.0002994 | 0 | 0.00027728
| 39.6 | 39.55 | 0.00029713 | 0.00024921 | 0.00018963 | 37.7 | 37.65 | 1572323.6999 | 1573598.0042 | 1570283.1258 |  |  |  | 37.7 | 37.65 | 1947252.795 |  | 0.00035401 | 0.0002968 | 0 | 0.00027516
| 39.7 | 39.65 | 0.00029001 | 0.00024221 | 0.00018348 | 37.8 | 37.75 | 1558547.0026 | 1559802.582 | 1556536.4004 |  |  |  | 37.8 | 37.75 | 1932202.3248 |  | 0.00035091 | 0.00029419 | 0 | 0.00027303
| 39.8 | 39.75 | 0.00028803 | 0.00024164 | 0.00018398 | 37.9 | 37.85 | 1544735.0589 | 1545972.0834 | 1542754.157 |  |  |  | 37.9 | 37.85 | 1917171.5358 |  | 0.00034781 | 0.00029159 | 0 | 0.00027091
| 39.9 | 39.85 | 0.00028112 | 0.00023559 | 0.00017868 | 38 | 37.95 | 1530836.6351 | 1532055.232 | 1528885.2302 |  |  |  | 38 | 37.95 | 1902173.69 |  | 0.00034468 | 0.00028896 | 0 | 0.00026879
| 40 | 39.95 | 0.00028041 | 0.00023455 | 0.00017704 | 38.1 | 38.05 | 1516967.0391 | 1518167.4255 | 1515044.7836 |  |  |  | 38.1 | 38.05 | 1887199.7133 |  | 0.00034156 | 0.00028634 | 0 | 0.00026667
| 40.1 | 40.05 | 0.00027123 | 0.00022675 | 0.00017098 | 38.2 | 38.15 | 1503072.8614 | 1504255.21 | 1501179.4796 |  |  |  | 38.2 | 38.15 | 1872160.3549 |  | 0.00033843 | 0.00028372 | 0 | 0.00026454
| 40.2 | 40.15 | 0.00027311 | 0.00022881 | 0.000173 | 38.3 | 38.25 | 1489220.0297 | 1490384.5624 | 1487355.1665 |  |  |  | 38.3 | 38.25 | 1857114.3064 |  | 0.00033532 | 0.0002811 | 0 | 0.00026242
| 40.3 | 40.25 | 0.00026344 | 0.00022017 | 0.00016598 | 38.4 | 38.35 | 1475301.5591 | 1476448.412 | 1473464.9971 |  |  |  | 38.4 | 38.35 | 1842083.9763 |  | 0.00033218 | 0.00027847 | 0 | 0.00026029
| 40.4 | 40.35 | 0.00026688 | 0.00022307 | 0.0001683 | 38.5 | 38.45 | 1461344.706 | 1462474.0347 | 1459536.1964 |  |  |  | 38.5 | 38.45 | 1827095.5195 |  | 0.00032904 | 0.00027584 | 0 | 0.00025818
| 40.5 | 40.45 | 0.00025767 | 0.00021631 | 0.00016311 | 38.6 | 38.55 | 1447395.23 | 1448507.2235 | 1445614.4704 |  |  |  | 38.6 | 38.55 | 1812123.9238 |  | 0.00032591 | 0.0002732 | 0 | 0.00025606
| 40.6 | 40.55 | 0.00025673 | 0.00021453 | 0.00016195 | 38.7 | 38.65 | 1433434.5858 | 1434529.4167 | 1431681.3002 |  |  |  | 38.7 | 38.65 | 1797087.2124 |  | 0.00032277 | 0.00027057 | 0 | 0.00025394
| 40.7 | 40.65 | 0.00025153 | 0.00021037 | 0.00015934 | 38.8 | 38.75 | 1419496.9659 | 1420574.8308 | 1417770.8401 |  |  |  | 38.8 | 38.75 | 1782042.8683 |  | 0.00031963 | 0.00026793 | 0 | 0.00025181
| 40.8 | 40.75 | 0.00024913 | 0.00020841 | 0.00015765 | 38.9 | 38.85 | 1405493.3944 | 1406554.4202 | 1403794.2256 |  |  |  | 38.9 | 38.85 | 1767020.0911 |  | 0.00031648 | 0.00026529 | 0 | 0.00024969
| 40.9 | 40.85 | 0.00024196 | 0.00020248 | 0.00015246 | 39 | 38.95 | 1391464.8364 | 1392509.1794 | 1389792.3746 |  |  |  | 39 | 38.95 | 1752027.7435 |  | 0.00031332 | 0.00026264 | 0 | 0.00024757
| 41 | 40.95 | 0.000245 | 0.0002045 | 0.00015429 | 39.1 | 39.05 | 1377427.9716 | 1378455.7988 | 1375781.9497 |  |  |  | 39.1 | 39.05 | 1737050.4452 |  | 0.00031016 | 0.00025999 | 0 | 0.00024545
| 41.1 | 41.05 | 0.00023494 | 0.00019636 | 0.00014776 | 39.2 | 39.15 | 1363393.5115 | 1364404.9959 | 1361773.6528 |  |  |  | 39.2 | 39.15 | 1722003.048 |  | 0.000307 | 0.00025734 | 0 | 0.00024333
| 41.2 | 41.15 | 0.00023709 | 0.00019776 | 0.00015002 | 39.3 | 39.25 | 1349355.3652 | 1350350.6732 | 1347761.4037 |  |  |  | 39.3 | 39.25 | 1706959.1593 |  | 0.00030385 | 0.00025469 | 0 | 0.0002412
| 41.3 | 41.25 | 0.00022828 | 0.00019151 | 0.00014388 | 39.4 | 39.35 | 1335247.0108 | 1336226.258 | 1333678.7616 |  |  |  | 39.4 | 39.35 | 1691943.323 |  | 0.00030067 | 0.00025203 | 0 | 0.00023908
| 41.4 | 41.35 | 0.00022813 | 0.00019071 | 0.00014367 | 39.5 | 39.45 | 1321131.7213 | 1322095.0684 | 1319588.9274 |  |  |  | 39.5 | 39.45 | 1676952.9848 |  | 0.00029749 | 0.00024936 | 0 | 0.00023696
| 41.5 | 41.45 | 0.0002215 | 0.0001854 | 0.00014022 | 39.6 | 39.55 | 1306999.1831 | 1307946.7814 | 1305481.6025 |  |  |  | 39.6 | 39.55 | 1661975.4954 |  | 0.00029431 | 0.00024669 | 0 | 0.00023484
| 41.6 | 41.55 | 0.00022236 | 0.00018543 | 0.00013965 | 39.7 | 39.65 | 1292894.174 | 1293826.2054 | 1291401.5159 |  |  |  | 39.7 | 39.65 | 1646937.2731 |  | 0.00029114 | 0.00024403 | 0 | 0.00023272
| 41.7 | 41.65 | 0.00021527 | 0.0001801 | 0.00013579 | 39.8 | 39.75 | 1278773.3169 | 1279689.9298 | 1277305.3438 |  |  |  | 39.8 | 39.75 | 1631916.483 |  | 0.00028796 | 0.00024136 | 0 | 0.0002306
| 41.8 | 41.75 | 0.00021445 | 0.00017902 | 0.0001346 | 39.9 | 39.85 | 1264566.8381 | 1265468.1296 | 1263123.3948 |  |  |  | 39.9 | 39.85 | 1616917.019 |  | 0.00028476 | 0.00023868 | 0 | 0.00022848
| 41.9 | 41.85 | 0.00020825 | 0.00017373 | 0.00013075 | 40 | 39.95 | 1250381.9407 | 1251268.0832 | 1248962.7519 |  |  |  | 40 | 39.95 | 1601939.0949 |  | 0.00028157 | 0.000236 | 0 | 0.00022636
| 42 | 41.95 | 0.00020703 | 0.00017308 | 0.00012981 | 40.1 | 40.05 | 1236165.4224 | 1237036.5483 | 1234770.2758 |  |  |  | 40.1 | 40.05 | 1586974.6749 |  | 0.00027837 | 0.00023332 | 0 | 0.00022425
| 42.1 | 42.05 | 0.00019989 | 0.00016574 | 0.00012535 | 40.2 | 40.15 | 1221913.2312 | 1222769.7045 | 1220541.5445 |  |  |  | 40.2 | 40.15 | 1571877.8231 |  | 0.00027516 | 0.00023063 | 0 | 0.00022211
| 42.2 | 42.15 | 0.00019979 | 0.00016702 | 0.0001264 | 40.3 | 40.25 | 1207532.1251 | 1208374.2306 | 1206183.443 |  |  |  | 40.3 | 40.25 | 1556710.8627 |  | 0.00027193 | 0.00022791 | 0 | 0.00021997
| 42.3 | 42.25 | 0.00019378 | 0.00016242 | 0.0001217 | 40.4 | 40.35 | 1193105.9234 | 1193933.7639 | 1191780.0816 |  |  |  | 40.4 | 40.35 | 1541560.4961 |  | 0.00026868 | 0.00022519 | 0 | 0.00021783
| 42.4 | 42.35 | 0.0001936 | 0.00016163 | 0.00012306 | 40.5 | 40.45 | 1178690.719 | 1179504.4349 | 1177387.4927 |  |  |  | 40.5 | 40.45 | 1526437.7084 |  | 0.00026544 | 0.00022247 | 0 | 0.00021569
| 42.5 | 42.45 | 0.00018672 | 0.00015552 | 0.00011723 | 40.6 | 40.55 | 1164258.5075 | 1165058.2185 | 1162977.7053 |  |  |  | 40.6 | 40.55 | 1511342.2188 |  | 0.00026219 | 0.00021974 | 0 | 0.00021356
| 42.6 | 42.55 | 0.00018487 | 0.00015542 | 0.00011852 | 40.7 | 40.65 | 1149872.0857 | 1150657.9532 | 1148613.4493 |  |  |  | 40.7 | 40.65 | 1496200.8215 |  | 0.00025895 | 0.00021703 | 0 | 0.00021142
| 42.7 | 42.65 | 0.0001785 | 0.00014949 | 0.00011194 | 40.8 | 40.75 | 1135429.8349 | 1136201.9493 | 1134193.2199 |  |  |  | 40.8 | 40.75 | 1481063.2854 |  | 0.0002557 | 0.0002143 | 0 | 0.00020928
| 42.8 | 42.75 | 0.0001794 | 0.00014966 | 0.00011257 | 40.9 | 40.85 | 1120962.4915 | 1121720.963 | 1119747.7214 |  |  |  | 40.9 | 40.85 | 1465933.0108 |  | 0.00025244 | 0.00021157 | 0 | 0.00020714
| 42.9 | 42.85 | 0.00017306 | 0.00014427 | 0.00010885 | 41 | 40.95 | 1106478.9764 | 1107223.9203 | 1105285.8669 |  |  |  | 41 | 40.95 | 1450835.1942 |  | 0.00024918 | 0.00020883 | 0 | 0.00020501
| 43 | 42.95 | 0.00017097 | 0.00014327 | 0.00010759 | 41.1 | 41.05 | 1092016.8314 | 1092748.3868 | 1090845.1599 |  |  |  | 41.1 | 41.05 | 1435757.5875 |  | 0.00024593 | 0.0002061 | 0 | 0.00020288
| 43.1 | 43.05 | 0.00016424 | 0.00013728 | 0.00010343 | 41.2 | 41.15 | 1077573.134 | 1078291.4366 | 1076422.6835 |  |  |  | 41.2 | 41.15 | 1420632.9894 |  | 0.00024267 | 0.00020338 | 0 | 0.00020074
| 43.2 | 43.15 | 0.00016621 | 0.00013825 | 0.00010403 | 41.3 | 41.25 | 1063066.0863 | 1063771.216 | 1061936.729 |  |  |  | 41.3 | 41.25 | 1405507.496 |  | 0.00023941 | 0.00020064 | 0 | 0.0001986
| 43.3 | 43.25 | 0.0001572 | 0.00013076 | 9.8387e-05 | 41.4 | 41.35 | 1048540.7263 | 1049232.792 | 1047432.288 |  |  |  | 41.4 | 41.35 | 1390389.8339 |  | 0.00023614 | 0.0001979 | 0 | 0.00019647
| 43.4 | 43.35 | 0.00015545 | 0.00013043 | 9.7447e-05 | 41.5 | 41.45 | 1033981.3818 | 1034660.481 | 1032893.7066 |  |  |  | 41.5 | 41.45 | 1375317.8554 |  | 0.00023286 | 0.00019515 | 0 | 0.00019434
| 43.5 | 43.45 | 0.00015139 | 0.00012565 | 9.4873e-05 | 41.6 | 41.55 | 1019477.3302 | 1020143.6178 | 1018410.1702 |  |  |  | 41.6 | 41.55 | 1360270.8708 |  | 0.0002296 | 0.00019241 | 0 | 0.00019221
| 43.6 | 43.55 | 0.00015111 | 0.00012637 | 9.4682e-05 | 41.7 | 41.65 | 1004965.8692 | 1005619.4575 | 1003919.0447 |  |  |  | 41.7 | 41.65 | 1345175.7089 |  | 0.00022633 | 0.00018967 | 0 | 0.00019008
| 43.7 | 43.65 | 0.00014333 | 0.00011903 | 8.9435e-05 | 41.8 | 41.75 | 990379.0198 | 991019.9759 | 989352.4234 |  |  |  | 41.8 | 41.75 | 1330078.0178 |  | 0.00022304 | 0.00018692 | 0 | 0.00018795
| 43.8 | 43.75 | 0.00014151 | 0.0001178 | 8.8316e-05 | 41.9 | 41.85 | 975777.1983 | 976405.6278 | 974770.6613 |  |  |  | 41.9 | 41.85 | 1314988.6442 |  | 0.00021976 | 0.00018416 | 0 | 0.00018581
| 43.9 | 43.85 | 0.00013601 | 0.00011337 | 8.5746e-05 | 42 | 41.95 | 961138.3885 | 961754.3817 | 960151.7663 |  |  |  | 42 | 41.95 | 1299948.4173 |  | 0.00021646 | 0.0001814 | 0 | 0.00018369
| 44 | 43.95 | 0.000134 | 0.00011163 | 8.3824e-05 | 42.1 | 42.05 | 946584.5128 | 947188.237 | 945617.5377 |  |  |  | 42.1 | 42.05 | 1284939.432 |  | 0.00021318 | 0.00017865 | 0 | 0.00018157
| 44.1 | 44.05 | 0.00012643 | 0.0001057 | 7.939e-05 | 42.2 | 42.15 | 931999.0411 | 932590.5874 | 931051.5671 |  |  |  | 42.2 | 42.15 | 1269877.9014 |  | 0.0002099 | 0.0001759 | 0 | 0.00017944
| 44.2 | 44.15 | 0.00012613 | 0.00010537 | 7.8863e-05 | 42.3 | 42.25 | 917322.1131 | 917901.5343 | 916394.0562 |  |  |  | 42.3 | 42.25 | 1254813.7759 |  | 0.0002066 | 0.00017313 | 0 | 0.00017731
| 44.3 | 44.25 | 0.00011916 | 9.9317e-05 | 7.4464e-05 | 42.4 | 42.35 | 902626.1192 | 903193.5131 | 901717.323 |  |  |  | 42.4 | 42.35 | 1239738.211 |  | 0.00020329 | 0.00017035 | 0 | 0.00017518
| 44.4 | 44.35 | 0.00011848 | 9.8461e-05 | 7.4011e-05 | 42.5 | 42.45 | 887920.4784 | 888475.9475 | 887030.7785 |  |  |  | 42.5 | 42.45 | 1224722.1783 |  | 0.00019998 | 0.00016758 | 0 | 0.00017306
| 44.5 | 44.45 | 0.00011057 | 9.2074e-05 | 6.8884e-05 | 42.6 | 42.55 | 873274.89 | 873818.5794 | 872404.0544 |  |  |  | 42.6 | 42.55 | 1209728.8903 |  | 0.00019668 | 0.00016481 | 0 | 0.00017094
| 44.6 | 44.55 | 0.00010938 | 9.1239e-05 | 6.8377e-05 | 42.7 | 42.65 | 858573.1382 | 859105.1193 | 857721.0527 |  |  |  | 42.7 | 42.65 | 1194681.5714 |  | 0.00019337 | 0.00016204 | 0 | 0.00016881
| 44.7 | 44.65 | 0.00010284 | 8.5644e-05 | 6.4472e-05 | 42.8 | 42.75 | 843806.7035 | 844327.0417 | 842973.2634 |  |  |  | 42.8 | 42.75 | 1179625.9559 |  | 0.00019004 | 0.00015925 | 0 | 0.00016669
| 44.8 | 44.75 | 0.00010138 | 8.4568e-05 | 6.3147e-05 | 42.9 | 42.85 | 828979.9346 | 829488.6975 | 828165.0318 |  |  |  | 42.9 | 42.85 | 1164548.7854 |  | 0.00018671 | 0.00015645 | 0 | 0.00016456
| 44.9 | 44.85 | 9.4117e-05 | 7.8368e-05 | 5.8858e-05 | 43 | 42.95 | 814183.0884 | 814680.3984 | 813386.5273 |  |  |  | 43 | 42.95 | 1149548.7078 |  | 0.00018337 | 0.00015366 | 0 | 0.00016244
| 45 | 44.95 | 9.3619e-05 | 7.7863e-05 | 5.8603e-05 | 43.1 | 43.05 | 799409.2833 | 799895.2568 | 798630.8772 |  |  |  | 43.1 | 43.05 | 1134567.7475 |  | 0.00018005 | 0.00015087 | 0 | 0.00016032
| 45.1 | 45.05 | 8.6637e-05 | 7.2386e-05 | 5.4365e-05 | 43.2 | 43.15 | 784561.7539 | 785036.4479 | 783801.4118 |  |  |  | 43.2 | 43.15 | 1119531.7475 |  | 0.0001767 | 0.00014807 | 0 | 0.0001582
| 45.2 | 45.15 | 8.4756e-05 | 7.0472e-05 | 5.2991e-05 | 43.3 | 43.25 | 769638.6606 | 770102.1304 | 768896.2942 |  |  |  | 43.3 | 43.25 | 1104476.2644 |  | 0.00017334 | 0.00014525 | 0 | 0.00015607
| 45.3 | 45.25 | 7.6957e-05 | 6.4435e-05 | 4.7893e-05 | 43.4 | 43.35 | 754653.355 | 755105.6636 | 753928.8636 |  |  |  | 43.4 | 43.35 | 1089393.6845 |  | 0.00016997 | 0.00014242 | 0 | 0.00015394
| 45.4 | 45.35 | 7.6858e-05 | 6.3674e-05 | 4.7625e-05 | 43.5 | 43.45 | 739731.5526 | 740172.8373 | 739024.7162 |  |  |  | 43.5 | 43.45 | 1074397.117 |  | 0.00016661 | 0.0001396 | 0 | 0.00015182
| 45.5 | 45.45 | 6.7856e-05 | 5.6455e-05 | 4.2342e-05 | 43.6 | 43.55 | 724794.8271 | 725225.1773 | 724105.5028 |  |  |  | 43.6 | 43.55 | 1059416.1972 |  | 0.00016325 | 0.00013678 | 0 | 0.0001497
| 45.6 | 45.55 | 6.6998e-05 | 5.5802e-05 | 4.1958e-05 | 43.7 | 43.65 | 709764.3165 | 710183.774 | 709092.4375 |  |  |  | 43.7 | 43.65 | 1044380.4445 |  | 0.00015986 | 0.00013395 | 0 | 0.00014758
| 45.7 | 45.65 | 5.9241e-05 | 4.9621e-05 | 3.7052e-05 | 43.8 | 43.75 | 694634.2496 | 695042.8526 | 693979.7548 |  |  |  | 43.8 | 43.75 | 1029316.0563 |  | 0.00015645 | 0.00013109 | 0 | 0.00014545
| 45.8 | 45.75 | 5.9572e-05 | 4.9627e-05 | 3.7189e-05 | 43.9 | 43.85 | 679452.1352 | 679849.9494 | 678814.9194 |  |  |  | 43.9 | 43.85 | 1014219.2879 |  | 0.00015304 | 0.00012823 | 0 | 0.00014331
| 45.9 | 45.85 | 5.1135e-05 | 4.2582e-05 | 3.1985e-05 | 44 | 43.95 | 664362.9327 | 664750.1079 | 663742.756 |  |  |  | 44 | 43.95 | 999218.0785 |  | 0.00014964 | 0.00012538 | 0 | 0.00014119
| 46 | 45.95 | 5.0913e-05 | 4.2255e-05 | 3.1825e-05 | 44.1 | 44.05 | 649227.0197 | 649603.6231 | 648623.775 |  |  |  | 44.1 | 44.05 | 984229.7074 |  | 0.00014623 | 0.00012252 | 0 | 0.00013908
| 46.1 | 46.05 | 4.3907e-05 | 3.6573e-05 | 2.7324e-05 | 44.2 | 44.15 | 633941.6654 | 634307.704 | 633355.3414 |  |  |  | 44.2 | 44.15 | 969185.3029 |  | 0.00014279 | 0.00011964 | 0 | 0.00013695
| 46.2 | 46.15 | 4.4712e-05 | 3.7303e-05 | 2.8103e-05 | 44.3 | 44.25 | 618558.6842 | 618914.1952 | 617989.2214 |  |  |  | 44.3 | 44.25 | 954098.464 |  | 0.00013932 | 0.00011673 | 0 | 0.00013482
| 46.3 | 46.25 | 3.5813e-05 | 2.9801e-05 | 2.2428e-05 | 44.4 | 44.35 | 603164.2851 | 603509.3547 | 602611.5454 |  |  |  | 44.4 | 44.35 | 938975.3261 |  | 0.00013586 | 0.00011383 | 0 | 0.00013268
| 46.4 | 46.35 | 3.7532e-05 | 3.1325e-05 | 2.3425e-05 | 44.5 | 44.45 | 587832.4406 | 588167.1966 | 587296.2199 |  |  |  | 44.5 | 44.45 | 923953.0299 |  | 0.0001324 | 0.00011093 | 0 | 0.00013056
| 46.5 | 46.45 | 2.9794e-05 | 2.4726e-05 | 1.8629e-05 | 44.6 | 44.55 | 572417.5692 | 572742.0551 | 571897.7976 |  |  |  | 44.6 | 44.55 | 908939.6014 |  | 0.00012893 | 0.00010802 | 0 | 0.00012844
| 46.6 | 46.55 | 2.9088e-05 | 2.4254e-05 | 1.8164e-05 | 44.7 | 44.65 | 556847.3 | 557161.5183 | 556343.9736 |  |  |  | 44.7 | 44.65 | 893862.5364 |  | 0.00012542 | 0.00010509 | 0 | 0.00012631
| 46.7 | 46.65 | 2.2786e-05 | 1.9093e-05 | 1.4249e-05 | 44.8 | 44.75 | 541189.209 | 541493.2005 | 540702.2625 |  |  |  | 44.8 | 44.75 | 878726.4252 |  | 0.0001219 | 0.00010213 | 0 | 0.00012417
| 46.8 | 46.75 | 2.3064e-05 | 1.9064e-05 | 1.4276e-05 | 44.9 | 44.85 | 525551.9776 | 525845.8437 | 525081.2489 |  |  |  | 44.9 | 44.85 | 863547.5885 |  | 0.00011838 | 9.918e-05 | 0 | 0.00012202
| 46.9 | 46.85 | 1.7055e-05 | 1.4231e-05 | 1.069e-05 | 45 | 44.95 | 509927.7067 | 510211.5432 | 509473.0423 |  |  |  | 45 | 44.95 | 848471.7976 |  | 0.00011486 | 9.6231e-05 | 0 | 0.00011989
| 47 | 46.95 | 1.699e-05 | 1.4126e-05 | 1.0552e-05 | 45.1 | 45.05 | 494203.9658 | 494477.8054 | 493765.3137 |  |  |  | 45.1 | 45.05 | 833397.6035 |  | 0.00011132 | 9.3263e-05 | 0 | 0.00011776
| 47.1 | 47.05 | 1.0445e-05 | 8.6917e-06 | 6.5172e-06 | 45.2 | 45.15 | 478288.1558 | 478552.0436 | 477865.4437 |  |  |  | 45.2 | 45.15 | 818221.0868 |  | 0.00010773 | 9.026e-05 | 0 | 0.00011562
| 47.2 | 47.15 | 1.0959e-05 | 9.1352e-06 | 6.8343e-06 | 45.3 | 45.25 | 462280.6082 | 462534.6376 | 461873.6866 |  |  |  | 45.3 | 45.25 | 802931.0377 |  | 0.00010413 | 8.7239e-05 | 0 | 0.00011346
| 47.3 | 47.25 | 5.8252e-06 | 4.8557e-06 | 3.6566e-06 | 45.4 | 45.35 | 446332.9377 | 446577.2188 | 445941.6305 |  |  |  | 45.4 | 45.35 | 787604.4205 |  | 0.00010054 | 8.4229e-05 | 0 | 0.00011129
| 47.4 | 47.35 | 5.6419e-06 | 4.6922e-06 | 3.5061e-06 | 45.5 | 45.45 | 430354.7139 | 430589.3064 | 429978.9258 |  |  |  | 45.5 | 45.45 | 772390.6118 |  | 9.6936e-05 | 8.1213e-05 | 0 | 0.00010914
| 47.5 | 47.45 | 1.5791e-06 | 1.3162e-06 | 9.8177e-07 | 45.6 | 45.55 | 414262.0002 | 414486.9175 | 413901.7095 |  |  |  | 45.6 | 45.55 | 757172.8318 |  | 9.3312e-05 | 7.8176e-05 | 0 | 0.00010699
| 47.6 | 47.55 | 1.4646e-06 | 1.223e-06 | 9.1604e-07 | 45.7 | 45.65 | 398007.0639 | 398222.2938 | 397662.2902 |  |  |  | 45.7 | 45.65 | 741873.4608 |  | 8.9651e-05 | 7.5109e-05 | 0 | 0.00010483
| 47.7 | 47.65 | 0 | 0 | 0 | 45.8 | 45.75 | 381692.8098 | 381898.3959 | 381363.4834 |  |  |  | 45.8 | 45.75 | 726464.5685 |  | 8.5976e-05 | 7.203e-05 | 0 | 0.00010265
| 47.8 | 47.75 | 0 | 0 | 0 | 45.9 | 45.85 | 365477.4902 | 365673.5607 | 365163.4061 |  |  |  | 45.9 | 45.85 | 711018.4466 |  | 8.2324e-05 | 6.897e-05 | 0 | 0.00010047
| 47.9 | 47.85 | 0 | 0 | 0 | 46 | 45.95 | 349197.8687 | 349384.463 | 348898.9633 |  |  |  | 46 | 45.95 | 695689.7343 |  | 7.8657e-05 | 6.5897e-05 | 0 | 9.8304e-05
| 48 | 47.95 | 0 | 0 | 0 | 46.1 | 46.05 | 332772.1751 | 332949.2892 | 332488.4553 |  |  |  | 46.1 | 46.05 | 680341.4491 |  | 7.4958e-05 | 6.2798e-05 | 0 | 9.6135e-05
| 48.1 | 48.05 | 0 | 0 | 0 | 46.2 | 46.15 | 316223.979 | 316391.6215 | 315955.4313 |  |  |  | 46.2 | 46.15 | 664899.8054 |  | 7.123e-05 | 5.9675e-05 | 0 | 9.3953e-05
| 48.2 | 48.15 | 0 | 0 | 0 | 46.3 | 46.25 | 299636.1018 | 299794.325 | 299382.6422 |  |  |  | 46.3 | 46.25 | 649325.191 |  | 6.7494e-05 | 5.6544e-05 | 0 | 9.1753e-05
| 48.3 | 48.25 | 0 | 0 | 0 | 46.4 | 46.35 | 283130.2668 | 283279.1867 | 282891.7096 |  |  |  | 46.4 | 46.35 | 633714.1335 |  | 6.3776e-05 | 5.3429e-05 | 0 | 8.9547e-05
| 48.4 | 48.35 | 0 | 0 | 0 | 46.5 | 46.45 | 266566.2139 | 266705.8718 | 266342.4929 |  |  |  | 46.5 | 46.45 | 618224.4463 |  | 6.0046e-05 | 5.0303e-05 | 0 | 8.7358e-05
| 48.5 | 48.45 | 0 | 0 | 0 | 46.6 | 46.55 | 249901.5659 | 250031.9808 | 249692.651 |  |  |  | 46.6 | 46.55 | 602699.7681 |  | 5.6292e-05 | 4.7158e-05 | 0 | 8.5164e-05
| 48.6 | 48.55 | 0 | 0 | 0 | 46.7 | 46.65 | 233106.0833 | 233227.2584 | 232911.9694 |  |  |  | 46.7 | 46.65 | 587071.5918 |  | 5.2509e-05 | 4.3989e-05 | 0 | 8.2956e-05
| 48.7 | 48.65 | 0 | 0 | 0 | 46.8 | 46.75 | 216285.2231 | 216397.2163 | 216105.8174 |  |  |  | 46.8 | 46.75 | 571289.7873 |  | 4.872e-05 | 4.0815e-05 | 0 | 8.0726e-05
| 48.8 | 48.75 | 0 | 0 | 0 | 46.9 | 46.85 | 199494.3362 | 199597.2336 | 199329.5011 |  |  |  | 46.9 | 46.85 | 555475.934 |  | 4.4938e-05 | 3.7646e-05 | 0 | 7.8491e-05
| 48.9 | 48.85 | 0 | 0 | 0 | 47 | 46.95 | 182686.8267 | 182780.6894 | 182536.464 |  |  |  | 47 | 46.95 | 539782.7968 |  | 4.1152e-05 | 3.4474e-05 | 0 | 7.6274e-05
| 49 | 48.95 | 0 | 0 | 0 | 47.1 | 47.05 | 166058.2419 | 166143.2308 | 165922.0942 |  |  |  | 47.1 | 47.05 | 524046.0282 |  | 3.7406e-05 | 3.1336e-05 | 0 | 7.405e-05
| 49.1 | 49.05 | 0 | 0 | 0 | 47.2 | 47.15 | 149010.8979 | 149086.8674 | 148889.1985 |  |  |  | 47.2 | 47.15 | 508196.8586 |  | 3.3566e-05 | 2.8119e-05 | 0 | 7.1811e-05
| 49.2 | 49.15 | 0 | 0 | 0 | 47.3 | 47.25 | 131841.3588 | 131908.3158 | 131734.0966 |  |  |  | 47.3 | 47.25 | 492180.5483 |  | 2.9699e-05 | 2.4879e-05 | 0 | 6.9547e-05
| 49.3 | 49.25 | 0 | 0 | 0 | 47.4 | 47.35 | 114609.7905 | 114667.7724 | 114516.9058 |  |  |  | 47.4 | 47.35 | 476130.9782 |  | 2.5817e-05 | 2.1627e-05 | 0 | 6.728e-05
| 49.4 | 49.35 | 0 | 0 | 0 | 47.5 | 47.45 | 97989.0497 | 98038.4329 | 97909.9397 |  |  |  | 47.5 | 47.45 | 460207.9986 |  | 2.2073e-05 | 1.8491e-05 | 0 | 6.503e-05
| 49.5 | 49.45 | 0 | 0 | 0 | 47.6 | 47.55 | 79072.8949 | 79112.5924 | 79009.3009 |  |  |  | 47.6 | 47.55 | 444232.2009 |  | 1.7812e-05 | 1.4921e-05 | 0 | 6.2772e-05
| 49.6 | 49.55 | 0 | 0 | 0 | 47.7 | 47.65 | 58683.6577 | 58713.0064 | 58636.6419 |  |  |  | 47.7 | 47.65 | 428133.2387 |  | 1.3219e-05 | 1.1074e-05 | 0 | 6.0497e-05
| 49.7 | 49.65 | 0 | 0 | 0 | 47.8 | 47.75 | 36509.3951 | 36527.5845 | 36480.2562 |  |  |  | 47.8 | 47.75 | 411870.4355 |  | 8.2243e-06 | 6.8895e-06 | 0 | 5.8199e-05
| 49.8 | 49.75 | 0 | 0 | 0 | 47.9 | 47.85 | 12577.5744 | 12583.8168 | 12567.5741 |  |  |  | 47.9 | 47.85 | 395565.7144 |  | 2.8333e-06 | 2.3734e-06 | 0 | 5.5895e-05
| 49.9 | 49.85 | 0 | 0 | 0 | [8 empty] | 48 | 47.95 | 379395.4722 |  |  |  | 0 | 5.361e-05
| 50 | 49.95 | 0 | 0 | 0 | [8 empty] | 48.1 | 48.05 | 363166.7834 |  |  |  | 0 | 5.1317e-05
| 50.1 | 50.05 | 0 | 0 | 0 | [8 empty] | 48.2 | 48.15 | 346821.1327 |  |  |  | 0 | 4.9007e-05
| 50.2 | 50.15 | 0 | 0 | 0 | [8 empty] | 48.3 | 48.25 | 330319.9207 |  |  |  | 0 | 4.6676e-05
| 50.3 | 50.25 | 0 | 0 | 0 | [8 empty] | 48.4 | 48.35 | 313761.7536 |  |  |  | 0 | 4.4336e-05
| 50.4 | 50.35 | 0 | 0 | 0 | [8 empty] | 48.5 | 48.45 | 297346.8079 |  |  |  | 0 | 4.2016e-05
| 50.5 | 50.45 | 0 | 0 | 0 | [8 empty] | 48.6 | 48.55 | 280874.9941 |  |  |  | 0 | 3.9689e-05
| 50.6 | 50.55 | 0 | 0 | 0 | [8 empty] | 48.7 | 48.65 | 264348.9227 |  |  |  | 0 | 3.7354e-05
| 50.7 | 50.65 | 0 | 0 | 0 | [8 empty] | 48.8 | 48.75 | 247685.2069 |  |  |  | 0 | 3.4999e-05
| 50.8 | 50.75 | 0 | 0 | 0 | [8 empty] | 48.9 | 48.85 | 230913.4055 |  |  |  | 0 | 3.2629e-05
| 50.9 | 50.85 | 0 | 0 | 0 | [8 empty] | 49 | 48.95 | 214278.5477 |  |  |  | 0 | 3.0279e-05
| 51 | 50.95 | 0 | 0 | 0 | [8 empty] | 49.1 | 49.05 | 197593.7879 |  |  |  | 0 | 2.7921e-05
[14 empty] | 49.2 | 49.15 | 181087.3098 |  |  |  | 0 | 2.5588e-05
[14 empty] | 49.3 | 49.25 | 164407.4518 |  |  |  | 0 | 2.3232e-05
[14 empty] | 49.4 | 49.35 | 147384.8251 |  |  |  | 0 | 2.0826e-05
[14 empty] | 49.5 | 49.45 | 130387.453 |  | 0 | 0 | 0 | 1.8424e-05
[14 empty] | 49.6 | 49.55 | 113312.0581 |  | 0 | 0 | 0 | 1.6012e-05
[14 empty] | 49.7 | 49.65 | 96891.4056 |  | 0 | 0 | 0 | 1.3691e-05
[14 empty] | 49.8 | 49.75 | 77803.7395 |  | 0 | 0 | 0 | 1.0994e-05
[14 empty] | 49.9 | 49.85 | 57527.6923 |  | 0 | 0 | 0 | 8.1289e-06
[14 empty] | 50 | 49.95 | 35682.9216 |  | 0 | 0 | 0 | 5.0422e-06
[14 empty] | 50.1 | 50.05 | 12257.087 |  | 0 | 0 | 0 | 1.732e-06
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
[18 empty] | 0 | 0 | 0 | 0
air
water

## EXP, air, 25 cm
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
3 June 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Papillon50 |  |  | The eBT-equivalent spectrum obtained from this measured spectrum is presented in the sheet "MC, water, 1 cm, eBT-eq."
Manufacturer: | Ariane Medical Systems
Spectra information
origin | measurement
detector | Ametek X-123 CdTe X-ray spectrometer with tungsten collimators with 50 um nominal diameter of aperture
data evaluation | deconvolution of fluence spectra using GRAVEL deconvolution algorithm and Monte Carlo simulated response matrix
authors | Czech Metrology Institute and Aarhus University
location | measured at Radiotherapy Department of Skejby Hospital in Aarhus, Aarhus, Denmark
publication | unpublished
DOI | N/A
year of the result | 2020
material | air
source-to-detector distance (cm) | 25
description of source-to-detector distance | on device axis, distance from the entrance window of the device
applicator | diameter 22 mm, 25 mm, 30 mm
other information | -
Device setting
nominal bias voltage (kV) | 50 | nominal, but the real voltage was obviously lower, see plot
bias voltage sigma (kV) | N/A
current (mA) | 2.7
characteristics of eBT spectra
Filtration material and density (g/cm3) | Al |  | 2.7 | 2.7 | 2.7 |  | 2.7 | 2.7 | 2.7
Emean (in F; keV) |  |  | 27.4372 | 27.4567 | 27.4885 |  | 27.3419 | 27.3634 | 27.4005 | average photon energy
Emean (in Kair; keV) |  |  | 22.1586 | 22.1871 | 22.2977 |  | 21.9878 | 22.0198 | 22.1441 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 20.19 | 20.2216 | 20.3448 |  | 20.0056 | 20.0412 | 20.1804 | energy of monoenergetic photons with the same HVL as the given spectrum
1st HVL in Al (in Kair; mm) |  |  | 0.76699 | 0.77045 | 0.7841 |  | 0.74688 | 0.75072 | 0.76592 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value
Measurement position to applicator end distance (cm): | [6 empty] | 23.5 | 23.14 | 22.54
recalculation of fluence spectrum from 25 cm to 0 cm
applicator diameter (mm) |  |  | 22 | 25 | 30 |  | 22 | 25 | 30
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Fluence (relative) | Fluence (relative) | Energy bin centre (keV) | Fluence (relative) | Fluence (relative) | Fluence (relative)
5.2
| 5.4 | 5.3 | 0 | 0 | 0 | 5.3 | 0 | 0 | 0
| 5.6 | 5.5 | 0 | 0 | 0 | 5.5 | 0 | 0 | 0
| 5.8 | 5.7 | 0 | 0 | 0 | 5.7 | 0 | 0 | 0 |  | APP22-DEG+00.0h-CUR2.7-DIS025-COL0.05-M120-04-08_14-54-18_GRAVEL.txt
| 6 | 5.9 | 0 | 0 | 0 | 5.9 | 0 | 0 | 0 |  | APP25-DEG+00.0h-CUR2.7-DIS025-COL0.05-M120-04-08_16-27-57_GRAVEL.txt
| 6.2 | 6.1 | 0 | 0 | 0 | 6.1 | 0 | 0 | 0 |  | APP30-DEG+00.0h-CUR2.7-DIS025-COL0.05-M120-04-08_15-02-57_GRAVEL.txt
| 6.4 | 6.3 | 0 | 0 | 0 | 6.3 | 0 | 0 | 0
| 6.6 | 6.5 | 0 | 0 | 0 | 6.5 | 0 | 0 | 0
| 6.8 | 6.7 | 0 | 0 | 0 | 6.7 | 0 | 0 | 0
| 7 | 6.9 | 0 | 0 | 0 | 6.9 | 0 | 0 | 0
| 7.2 | 7.1 | 0 | 0 | 0 | 7.1 | 0 | 0 | 0
| 7.4 | 7.3 | 0 | 0 | 0 | 7.3 | 0 | 0 | 0
| 7.6 | 7.5 | 0 | 0 | 0 | 7.5 | 0 | 0 | 0
| 7.8 | 7.7 | 0 | 0 | 0 | 7.7 | 0 | 0 | 0
| 8 | 7.9 | 0 | 0 | 0 | 7.9 | 0 | 0 | 0
| 8.2 | 8.1 | 0 | 0 | 0 | 8.1 | 0 | 0 | 0
| 8.4 | 8.3 | 0 | 0 | 0 | 8.3 | 0 | 0 | 0
| 8.6 | 8.5 | 0.0073114 | 0 | 0 | 8.5 | 0.0092437 | 0 | 0
| 8.8 | 8.7 | 0.0074342 | 0 | 0 | 8.7 | 0.0092528 | 0 | 0
| 9 | 8.9 | 0 | 0 | 0 | 8.9 | 0 | 0 | 0
| 9.2 | 9.1 | 0 | 0 | 0 | 9.1 | 0 | 0 | 0
| 9.4 | 9.3 | 0 | 0 | 0 | 9.3 | 0 | 0 | 0
| 9.6 | 9.5 | 0 | 0 | 0 | 9.5 | 0 | 0 | 0
| 9.8 | 9.7 | 0.0080302 | 3.4457e-07 | 15086.1 | 9.7 | 0.0094101 | 4.028e-07 | 17564.3948
| 10 | 9.9 | 166659000 | 164700000 | 165397000 | 9.9 | 193500838.8427 | 190789368.4039 | 190867660.4736
| 10.2 | 10.1 | 179910000 | 167206000 | 164353000 | 10.1 | 207110453.2575 | 192071026.5221 | 188116306.435
| 10.4 | 10.3 | 87690900 | 98495500 | 7602620 | 10.3 | 100154989.8095 | 112266520.1351 | 8636216.1444
| 10.6 | 10.5 | 9494040 | 9104030 | 8746730 | 10.5 | 10764645.4168 | 10302597.0131 | 9866565.9043
| 10.8 | 10.7 | 10964300 | 10189100 | 9815770 | 10.7 | 12347974.7169 | 11454073.1156 | 11000962.2603
| 11 | 10.9 | 12824200 | 11722000 | 11327100 | 10.9 | 14352434.9651 | 13096281.0857 | 12618758.0867
| 11.2 | 11.1 | 14859900 | 14052400 | 13777600 | 11.1 | 16534345.6713 | 15610300.4969 | 15263368.4528
| 11.4 | 11.3 | 17209000 | 16905700 | 16813400 | 11.3 | 19044995.4784 | 18680305.224 | 18530293.8631
| 11.6 | 11.5 | 109968000 | 102093000 | 102745000 | 11.5 | 121090104.1379 | 112252831.0973 | 112692164.1171
| 11.8 | 11.7 | 146892000 | 151432000 | 149422000 | 11.7 | 160992906.2521 | 165735835.0415 | 163153696.5443
| 12 | 11.9 | 83467900 | 97294100 | 103006000 | 11.9 | 91081694.901 | 106027213.5139 | 112001904.9682
| 12.2 | 12.1 | 52724100 | 51036900 | 50414200 | 12.1 | 57299302.3396 | 55395031.0208 | 54603021.4629
| 12.4 | 12.3 | 62732600 | 61118800 | 60595800 | 12.3 | 67916822.2568 | 66089219.382 | 65390985.4981
| 12.6 | 12.5 | 73379700 | 71734400 | 71231600 | 12.5 | 79160783.6817 | 77296014.2254 | 76605765.6362
| 12.8 | 12.7 | 82317100 | 79711500 | 79388000 | 12.7 | 88505880.9818 | 85609265.4298 | 85104172.3428
| 13 | 12.9 | 94870000 | 91922800 | 91703700 | 12.9 | 101683184.2064 | 98419706.4462 | 98011413.3817
| 13.2 | 13.1 | 107039000 | 104151000 | 104006000 | 13.1 | 114388881.9984 | 111189399.5684 | 110846490.9324
| 13.4 | 13.3 | 120106000 | 116989000 | 116928000 | 13.3 | 127998565.2562 | 124555239.0228 | 124288166.0579
| 13.6 | 13.5 | 134092000 | 131058000 | 131144000 | 13.5 | 142532289.5175 | 139177109.4479 | 139051552.1711
| 13.8 | 13.7 | 147123000 | 144236000 | 144420000 | 13.7 | 156000900.2462 | 152802473.0464 | 152768690.2483
| 14 | 13.9 | 160838000 | 158111000 | 158379000 | 13.9 | 170150294.5192 | 167121245.8862 | 167164122.355
| 14.2 | 14.1 | 175080000 | 172495000 | 172836000 | 14.1 | 184814093.1295 | 181934508.087 | 182042509.4899
| 14.4 | 14.3 | 189437000 | 186908000 | 187312000 | 14.3 | 199558786.4075 | 196737718.944 | 196901111.0283
| 14.6 | 14.5 | 202965000 | 200597000 | 201099000 | 14.5 | 213394890.6936 | 210743364.3657 | 211000623.5238
| 14.8 | 14.7 | 217052000 | 214750000 | 215334000 | 14.7 | 227787252.3495 | 225204788.7222 | 225539058.8608
| 15 | 14.9 | 231689000 | 229380000 | 230029000 | 14.9 | 242726190.3596 | 240135934.7191 | 240529398.8854
| 15.2 | 15.1 | 245428000 | 243242000 | 243984000 | 15.1 | 256696292.5609 | 254235035.439 | 254718461.8861
| 15.4 | 15.3 | 259244000 | 257193000 | 258031000 | 15.3 | 270722730.5871 | 268402716.5305 | 268979535.6949
| 15.6 | 15.5 | 274654000 | 272259000 | 273163000 | 15.5 | 286389884.8564 | 283710635.0168 | 284348724.8343
| 15.8 | 15.7 | 288618000 | 286282000 | 287270000 | 15.7 | 300527255.4425 | 297910275.8993 | 298629951.3687
| 16 | 15.9 | 303067000 | 300667000 | 301729000 | 15.9 | 315151002.8894 | 312468100.9697 | 313258918.3829
| 16.2 | 16.1 | 317549000 | 315030000 | 316156000 | 16.1 | 329791411.8902 | 326985755.3596 | 327837700.8434
| 16.4 | 16.3 | 332060000 | 329377000 | 330555000 | 16.3 | 344445786.8839 | 341471091.3148 | 342372075.8563
| 16.6 | 16.5 | 346548000 | 343655000 | 344902000 | 16.5 | 359061587.4799 | 355870687.2264 | 356838684.3745
| 16.8 | 16.7 | 361038000 | 357949000 | 359261000 | 16.7 | 373666082.6835 | 370273977.4784 | 371305088.7049
| 17 | 16.9 | 374261000 | 371499000 | 372937000 | 16.9 | 386948442.2625 | 383896700.3636 | 385054797.3772
| 17.2 | 17.1 | 389774000 | 386576000 | 388020000 | 17.1 | 402587588.5417 | 399086656.9308 | 400246710.4033
| 17.4 | 17.3 | 403446000 | 400473000 | 402033000 | 17.3 | 416314850.8728 | 413048292.5035 | 414324988.5194
| 17.6 | 17.5 | 415437000 | 413177000 | 414915000 | 17.5 | 428301379.1984 | 425772439.6883 | 427230639.7342
| 17.8 | 17.7 | 429366000 | 427304000 | 429195000 | 17.7 | 442280228.2416 | 439956437.3707 | 441569206.6376
| 18 | 17.9 | 441097000 | 439970000 | 442096000 | 17.9 | 453990031.3503 | 452630276.6695 | 454483018.8614
| 18.2 | 18.1 | 451393000 | 451476000 | 453838000 | 18.1 | 464221449.0939 | 464107526.6929 | 466201929.4275
| 18.4 | 18.3 | 461121000 | 463193000 | 466153000 | 18.3 | 473869125.9332 | 475799594.9302 | 478506868.948
| 18.6 | 18.5 | 470676000 | 474340000 | 477638000 | 18.5 | 483340116.2824 | 486904620.288 | 489957730.6497
| 18.8 | 18.7 | 477630000 | 482739000 | 486532000 | 18.7 | 490143269.9565 | 495189898.6972 | 498751299.2073
| 19 | 18.9 | 483731000 | 490001000 | 494230000 | 18.9 | 496076547.1506 | 502312605.7361 | 506321971.1659
| 19.2 | 19.1 | 489284000 | 496509000 | 501006000 | 19.1 | 501454023.6619 | 508667247.0938 | 512952496.7157
| 19.4 | 19.3 | 494442000 | 501937000 | 506485000 | 19.3 | 506433198.803 | 513921279.3592 | 518260693.4068
| 19.6 | 19.5 | 514500000 | 512972000 | 515915000 | 19.5 | 526671343.2044 | 524919146.9017 | 527615627.0079
| 19.8 | 19.7 | 527999000 | 531576000 | 534321000 | 19.7 | 540188262.9352 | 543657726.3259 | 546146769.9111
| 20 | 19.9 | 524716000 | 521101000 | 521525000 | 19.9 | 536542101.4821 | 532663727.1893 | 532793861.6526
| 20.2 | 20.1 | 510471000 | 513800000 | 517512000 | 20.1 | 521709012.2403 | 524936156.3264 | 528434726.1086
| 20.4 | 20.3 | 514190000 | 517009000 | 520671000 | 20.3 | 525252885.4818 | 527960341.5899 | 531411011.4084
| 20.6 | 20.5 | 517721000 | 519746000 | 523296000 | 20.5 | 528611269.5009 | 530509660.923 | 533849366.5588
| 20.8 | 20.7 | 522521000 | 524300000 | 527673000 | 20.7 | 533271103.0865 | 534919798.0433 | 538081270.0443
| 21 | 20.9 | 526384000 | 527856000 | 531131000 | 20.9 | 536980000.565 | 538317253.644 | 541381607.9935
| 21.2 | 21.1 | 531140000 | 532012000 | 535156000 | 21.1 | 541605009.0024 | 542332064.3088 | 545265354.2442
| 21.4 | 21.3 | 535628000 | 536190000 | 539202000 | 21.3 | 545961386.6583 | 546374268.6245 | 549175484.5477
| 21.6 | 21.5 | 539871000 | 539905000 | 542791000 | 21.5 | 550072697.5547 | 549949604.0223 | 552625098.7109
| 21.8 | 21.7 | 543965000 | 543498000 | 546296000 | 21.7 | 554036808.8987 | 553405606.4775 | 555994115.6489
| 22 | 21.9 | 586423000 | 590862000 | 595348000 | 21.9 | 597065644.0395 | 601419475.5982 | 605707420.8293
| 22.2 | 22.1 | 619723000 | 648758000 | 656673000 | 22.1 | 630750653.7297 | 660123927.4758 | 667876760.168
| 22.4 | 22.3 | 595714000 | 610693000 | 613661000 | 22.3 | 606111101.9845 | 621186858.6379 | 623930165.4345
| 22.6 | 22.5 | 557478000 | 555482000 | 558174000 | 22.5 | 567024197.473 | 564847080.7037 | 567338468.8749
| 22.8 | 22.7 | 560362000 | 558115000 | 560773000 | 22.7 | 569779507.5802 | 567349872.3132 | 569809332.4629
| 23 | 22.9 | 563120000 | 560621000 | 563244000 | 22.9 | 572411089.3194 | 569728012.9146 | 572154513.9121
| 23.2 | 23.1 | 565542000 | 562827000 | 565525000 | 23.1 | 574705472.5622 | 571805670.3437 | 574310978.2608
| 23.4 | 23.3 | 567354000 | 564522000 | 567338000 | 23.3 | 576384410.7086 | 573368613.9957 | 575996464.1447
| 23.6 | 23.5 | 570167000 | 567232000 | 570011000 | 23.5 | 579084431.4666 | 575966582.7108 | 578559087.8288
| 23.8 | 23.7 | 571763000 | 568825000 | 571779000 | 23.7 | 580552452.703 | 577434324.5037 | 580206996.9775
| 24 | 23.9 | 572435000 | 569711000 | 572965000 | 23.9 | 581086700.495 | 578188647.8016 | 581268401.663
| 24.2 | 24.1 | 573937000 | 571349000 | 574776000 | 24.1 | 582467769.1044 | 579710259.9967 | 582967764.2218
| 24.4 | 24.3 | 574291000 | 571981000 | 575737000 | 24.3 | 582687950.3385 | 580215140.39 | 583808805.6405
| 24.6 | 24.5 | 576814000 | 572251000 | 576350000 | 24.5 | 585112617.8363 | 580356958.8611 | 584300882.0455
| 24.8 | 24.7 | 583422000 | 572140000 | 576572000 | 24.7 | 591683251.8099 | 580116529.5886 | 584400483.5154
| 25 | 24.9 | 572551000 | 571240000 | 576019000 | 24.9 | 580532434.6711 | 579080337.0191 | 583718571.5978
| 25.2 | 25.1 | 569520000 | 587604000 | 590185000 | 25.1 | 577337864.5203 | 595545708.4749 | 597953410.5058
| 25.4 | 25.3 | 568404000 | 586311000 | 592080000 | 25.3 | 576089203.3356 | 594116073.4991 | 599756182.3384
| 25.6 | 25.5 | 565892000 | 564401000 | 569621000 | 25.5 | 573429992.0473 | 571803207.7992 | 576896729.1904
| 25.8 | 25.7 | 563324000 | 561828000 | 565858000 | 25.7 | 570718472.8008 | 569089132.442 | 572980402.9654
| 26 | 25.9 | 560633000 | 559350000 | 563128000 | 25.9 | 567886633.5819 | 566475465.4112 | 570114438.6911
| 26.2 | 26.1 | 557778000 | 556660000 | 560358000 | 26.1 | 564892849.6053 | 563651133.5622 | 567211987.382
| 26.4 | 26.3 | 554343000 | 553538000 | 557039000 | 26.3 | 561315812.3112 | 560393366.4159 | 563757772.7006
| 26.6 | 26.5 | 552056000 | 551348000 | 554834000 | 26.5 | 558905101.5514 | 558082892.481 | 561434699.9226
| 26.8 | 26.7 | 549232000 | 548623000 | 552066000 | 26.7 | 555954359.4478 | 555234421.5904 | 558545400.4228
| 27 | 26.9 | 545922000 | 545394000 | 548783000 | 26.9 | 552515328.0781 | 551879447.1425 | 555138563.8639
| 27.2 | 27.1 | 542424000 | 542037000 | 545365000 | 27.1 | 548889652.868 | 548398483.6888 | 551598636.6351
| 27.4 | 27.3 | 539005000 | 538672000 | 541961000 | 27.3 | 545347425.2702 | 544912846.0488 | 548076228.3716
| 27.6 | 27.5 | 535692000 | 535360000 | 538561000 | 27.5 | 541915790.926 | 541484106.6784 | 544561094.9584
| 27.8 | 27.7 | 532200000 | 531962000 | 535114000 | 27.7 | 538306308.0664 | 537971549.565 | 541001552.7972
| 28 | 27.9 | 528648000 | 528455000 | 531571000 | 27.9 | 534639281.9161 | 534351836.7911 | 537347973.0204
| 28.2 | 28.1 | 524994000 | 524823000 | 527865000 | 28.1 | 530872151.1437 | 530608723.2182 | 533532563.2565
| 28.4 | 28.3 | 521168000 | 521149000 | 524151000 | 28.3 | 526934070.7175 | 526826053.0876 | 529711923.3781
| 28.6 | 28.5 | 517541000 | 517643000 | 520602000 | 28.5 | 523200056.574 | 523215997.7976 | 526060766.4682
| 28.8 | 28.7 | 513750000 | 513891000 | 516803000 | 28.7 | 519303005.4831 | 519359987.6841 | 522159631.9661
| 29 | 28.9 | 510079000 | 510333000 | 513189000 | 28.9 | 515529914.5439 | 515702645.9611 | 518447972.0325
| 29.2 | 29.1 | 506214000 | 506434000 | 509206000 | 29.1 | 511563327.2951 | 511703244.3347 | 514366017.4797
| 29.4 | 29.3 | 502661000 | 502860000 | 505593000 | 29.3 | 507914509.3613 | 508034665.2115 | 510660211.5376
| 29.6 | 29.5 | 498988000 | 499132000 | 501801000 | 29.5 | 504146800.6283 | 504212837.1585 | 506775905.3455
| 29.8 | 29.7 | 495322000 | 495355000 | 497950000 | 29.7 | 500388457.9382 | 500343786.9479 | 502834253.3807
| 30 | 29.9 | 491575000 | 491485000 | 494013000 | 29.9 | 496550508.9482 | 496383012.9367 | 498807934.3996
| 30.2 | 30.1 | 487612000 | 487516000 | 490008000 | 30.1 | 492497279.1335 | 492325125.9121 | 494715774.4012
| 30.4 | 30.3 | 483869000 | 483636000 | 486080000 | 30.3 | 488669022.3135 | 488359856.3212 | 490704039.6906
| 30.6 | 30.5 | 480035000 | 479788000 | 482201000 | 30.5 | 484750798.7801 | 484428819.3568 | 486743654.0159
| 30.8 | 30.7 | 476112000 | 475834000 | 478242000 | 30.7 | 480744586.7968 | 480392617.3411 | 482704334.9911
| 31 | 30.9 | 471999000 | 471771000 | 474219000 | 30.9 | 476548372.6919 | 476248186.7678 | 478602189.1607
| 31.2 | 31.1 | 467915000 | 467705000 | 470136000 | 31.1 | 472383236.1385 | 472102491.3065 | 474441209.2937
| 31.4 | 31.3 | 463899000 | 463557000 | 465953000 | 31.3 | 468288477.4527 | 467875735.8859 | 470180989.2878
| 31.6 | 31.5 | 459721000 | 459270000 | 461632000 | 31.5 | 464031864.1352 | 463510357.4426 | 465783155.7484
| 31.8 | 31.7 | 455420000 | 454902000 | 457245000 | 31.7 | 459652743.3793 | 459064865.2892 | 461320328.9152
| 32 | 31.9 | 451240000 | 450622000 | 452935000 | 31.9 | 455397338.7864 | 454709757.8018 | 456936734.6058
| 32.2 | 32.1 | 446904000 | 446308000 | 448621000 | 32.1 | 450986038.7049 | 450321864.8775 | 452550594.5951
| 32.4 | 32.3 | 442385000 | 441607000 | 443780000 | 32.3 | 446391584.7386 | 445544996.6664 | 447634318.6705
| 32.6 | 32.5 | 437999000 | 437235000 | 439388000 | 32.5 | 441932807.5465 | 441101523.1497 | 443172380.5418
| 32.8 | 32.7 | 434193000 | 433268000 | 435379000 | 32.7 | 438060611.5106 | 437067991.1926 | 439098073.5991
| 33 | 32.9 | 430248000 | 429159000 | 431229000 | 32.9 | 434049473.1755 | 432892511.403 | 434882835.0173
| 33.2 | 33.1 | 425966000 | 424684000 | 426700000 | 33.1 | 429699645.6729 | 428349139.2902 | 430286652.7611
| 33.4 | 33.3 | 421927000 | 420484000 | 422450000 | 33.3 | 425596202.6143 | 424084397.8969 | 425973050.1325
| 33.6 | 33.5 | 417325000 | 415965000 | 417921000 | 33.5 | 420926100.1076 | 419499145.7232 | 421379316.3073
| 33.8 | 33.7 | 413566000 | 411593000 | 413421000 | 33.7 | 417107451.9023 | 415063336.5037 | 416815996.7236
| 34 | 33.9 | 408831000 | 407017000 | 408852000 | 33.9 | 412305597.5369 | 410422967.8017 | 412184250.8569
| 34.2 | 34.1 | 404840000 | 402858000 | 404631000 | 34.1 | 408255197.7817 | 406204200.3854 | 407904429.7452
| 34.4 | 34.3 | 399990000 | 398229000 | 400024000 | 34.3 | 403339655.058 | 401512609.8405 | 403236543.171
| 34.6 | 34.5 | 395309000 | 393884000 | 395728000 | 34.5 | 398595639.7421 | 397108420.3399 | 398883183.9475
| 34.8 | 34.7 | 390660000 | 389758000 | 391695000 | 34.7 | 393884957.054 | 392926021.6073 | 394795887.7182
| 35 | 34.9 | 385652000 | 384990000 | 386959000 | 34.9 | 388813364.1887 | 388097396.615 | 390000987.5039
| 35.2 | 35.1 | 380799000 | 380909000 | 383075000 | 35.1 | 383899075.6479 | 383962277.0355 | 386065710.2081
| 35.4 | 35.3 | 375842000 | 376454000 | 378732000 | 35.3 | 378880940.5708 | 379451074.0988 | 381668725.9188
| 35.6 | 35.5 | 370768000 | 371887000 | 374283000 | 35.5 | 373745842.1211 | 374827893.302 | 377165800.0578
| 35.8 | 35.7 | 365806000 | 367380000 | 369868000 | 35.7 | 368724597.3628 | 370266076.8631 | 372697994.4494
| 36 | 35.9 | 360688000 | 362403000 | 364892000 | 35.9 | 363547036.5125 | 365231453.1281 | 367665756.1084
| 36.2 | 36.1 | 355733000 | 357948000 | 360537000 | 36.1 | 358534669.0709 | 360723760.4489 | 363260070.5419
| 36.4 | 36.3 | 350571000 | 352867000 | 355441000 | 36.3 | 353314548.8494 | 355586050.5534 | 358108602.0471
| 36.6 | 36.5 | 345752000 | 348502000 | 351108000 | 36.5 | 348440958.7528 | 351170667.1085 | 353726649.6945
| 36.8 | 36.7 | 340694000 | 343568000 | 346163000 | 36.7 | 343327326.5209 | 346182705.5068 | 348728892.9692
| 37 | 36.9 | 335623000 | 338661000 | 341249000 | 36.9 | 338201399.222 | 341222731.6556 | 343763131.0423
| 37.2 | 37.1 | 330661000 | 333800000 | 336339000 | 37.1 | 333186087.6312 | 336309862.7107 | 338802140.5298
| 37.4 | 37.3 | 325330000 | 328439000 | 330949000 | 37.3 | 327799726.3666 | 330893990.1244 | 333358376.5422
| 37.6 | 37.5 | 320451000 | 323597000 | 326071000 | 37.5 | 322869539.7139 | 326001731.1439 | 328431059.9581
| 37.8 | 37.7 | 315076000 | 318158000 | 320601000 | 37.7 | 317440333.9883 | 320508752.6949 | 322908161.6595
| 38 | 37.9 | 309934000 | 313005000 | 315424000 | 37.9 | 312246592.4861 | 315304597.867 | 317681068.0319
| 38.2 | 38.1 | 304629000 | 307646000 | 310035000 | 38.1 | 306889327.8813 | 309893617.1693 | 312241131.0895
| 38.4 | 38.3 | 299139000 | 302057000 | 304402000 | 38.3 | 301346379.0722 | 304251642.5445 | 306556131.1377
| 38.6 | 38.5 | 293749000 | 296613000 | 298921000 | 38.5 | 295904841.4146 | 298756392.6339 | 301024865.3813
| 38.8 | 38.7 | 288347000 | 291223000 | 293515000 | 38.7 | 290451867.2159 | 293316178.2173 | 295569759.6141
| 39 | 38.9 | 282936000 | 286070000 | 288500000 | 38.9 | 284990462.1232 | 288115284.0978 | 290508988.9183
| 39.2 | 39.1 | 277370000 | 280332000 | 282689000 | 39.1 | 279373555.5457 | 282325820.6862 | 284647271.6572
| 39.4 | 39.3 | 271814000 | 274865000 | 277269000 | 39.3 | 273767334.0006 | 276809893.2754 | 279179858.1978
| 39.6 | 39.5 | 266272000 | 269228000 | 271591000 | 39.5 | 268175808.7107 | 271123351.6565 | 273453241.3288
| 39.8 | 39.7 | 260558000 | 263307000 | 265578000 | 39.7 | 262411639.1922 | 265151399.8558 | 267389907.3232
| 40 | 39.9 | 254974000 | 257612000 | 259839000 | 39.9 | 256778966.175 | 259407606.8739 | 261603009.4583
| 40.2 | 40.1 | 249226000 | 251620000 | 253744000 | 40.1 | 250981746.151 | 253365362.5317 | 255458304.1025
| 40.4 | 40.3 | 243629000 | 245804000 | 247846000 | 40.3 | 245337183.8239 | 247500941.1903 | 249512523.8607
| 40.6 | 40.5 | 237968000 | 239916000 | 241856000 | 40.5 | 239628691.559 | 241564549.3845 | 243474644.8604
| 40.8 | 40.7 | 232179000 | 233882000 | 235715000 | 40.7 | 233791817.4555 | 235481674.1108 | 237285269.0555
| 41 | 40.9 | 226467000 | 227924000 | 229644000 | 40.9 | 228032978.0038 | 229475827.0996 | 231166862.4772
| 41.2 | 41.1 | 220512000 | 221753000 | 223366000 | 41.1 | 222029950.2407 | 223256029.3356 | 224840577.1538
| 41.4 | 41.3 | 214613000 | 215810000 | 217331000 | 41.3 | 216083793.0802 | 217266263.1643 | 218759376.2719
| 41.6 | 41.5 | 208606000 | 209647000 | 211063000 | 41.5 | 210029370.0233 | 211055485.9466 | 212444111.6462
| 41.8 | 41.7 | 202574000 | 203410000 | 204722000 | 41.7 | 203950242.5278 | 204770681.6064 | 206055833.7657
| 42 | 41.9 | 196521000 | 197140000 | 198346000 | 41.9 | 197850428.0589 | 198453117.7365 | 199632783.546
| 42.2 | 42.1 | 190314000 | 190662000 | 191760000 | 42.1 | 191596021.0357 | 191926624.7945 | 192998821.8691
| 42.4 | 42.3 | 184216000 | 184334000 | 185327000 | 42.3 | 185451787.567 | 185551573.389 | 186519289.8444
| 42.6 | 42.5 | 177909000 | 177834000 | 178731000 | 42.5 | 179097583.1483 | 179003821.9474 | 179876139.6541
| 42.8 | 42.7 | 171310000 | 171255000 | 172069000 | 42.7 | 172449861.9637 | 172376982.8391 | 173166992.4123
| 43 | 42.9 | 165122000 | 164757000 | 165472000 | 42.9 | 166216295.7576 | 165832095.7532 | 166523675.3234
| 43.2 | 43.1 | 158546000 | 157853000 | 158496000 | 43.1 | 159592567.5 | 158878978.7761 | 159499362.5431
| 43.4 | 43.3 | 152226000 | 151224000 | 151775000 | 43.3 | 153226931.9368 | 152203061.7855 | 152732070.22
| 43.6 | 43.5 | 145403000 | 144336000 | 144860000 | 43.5 | 146355388.2856 | 145266870.1297 | 145769949.3789
| 43.8 | 43.7 | 138374000 | 137543000 | 138085000 | 43.7 | 139276902.7828 | 138426687.7255 | 138949094.6069
| 44 | 43.9 | 131242000 | 130578000 | 131122000 | 43.9 | 132095150.5076 | 131413789.173 | 131939441.8889
| 44.2 | 44.1 | 123464000 | 123386000 | 123970000 | 44.1 | 124263612.5969 | 124172826.8365 | 124739989.1556
| 44.4 | 44.3 | 116504000 | 116597000 | 117191000 | 44.3 | 117255772.4485 | 117337810.3111 | 117916218.3125
| 44.6 | 44.5 | 108993000 | 109482000 | 110143000 | 44.5 | 109693761.0291 | 110175087.7266 | 110822136.9746
| 44.8 | 44.7 | 101144000 | 102166000 | 102957000 | 44.7 | 101791972.2118 | 102810461.3456 | 103589559.6395
| 45 | 44.9 | 92636400 | 94182000 | 95076500 | 44.9 | 93227773.2592 | 94774000.7047 | 95658580.0854
| 45.2 | 45.1 | 84802100 | 86478400 | 87508400 | 45.1 | 85341571.9867 | 87020081.8851 | 88042277.6901
| 45.4 | 45.3 | 76476200 | 78363300 | 79512900 | 45.3 | 76961029.5493 | 78852458.8842 | 79996326.3366
| 45.6 | 45.5 | 67744300 | 69704200 | 70870500 | 45.5 | 68172310.0246 | 70137825.3348 | 71299914.5949
| 45.8 | 45.7 | 59762500 | 61443700 | 62671100 | 45.7 | 60138810.1667 | 61824651.0238 | 63049555.5019
| 46 | 45.9 | 51513800 | 52761700 | 53828000 | 45.9 | 51837091.3748 | 53087734.8093 | 54151973.3377
| 46.2 | 46.1 | 44009800 | 44705300 | 45696600 | 46.1 | 44285089.9121 | 44980643.3827 | 45970729.3071
| 46.4 | 46.3 | 36640100 | 36773400 | 37574200 | 46.3 | 36868546.5898 | 36999154.6008 | 37798871.8334
| 46.6 | 46.5 | 29504000 | 29152300 | 29683500 | 46.5 | 29687363.3501 | 29330693.6213 | 29860420.345
| 46.8 | 46.7 | 22961600 | 22390300 | 22761100 | 46.7 | 23103850.4335 | 22526879.7429 | 22896330.8902
| 47 | 46.9 | 16724400 | 16178500 | 16368400 | 46.9 | 16827685.1197 | 16276878.5587 | 16465344.8674
| 47.2 | 47.1 | 10987700 | 10771800 | 10824200 | 47.1 | 11055346.518 | 10837098.3172 | 10888109.587
| 47.4 | 47.3 | 5937590 | 6416990 | 6425040 | 47.3 | 5974033.1349 | 6455770.3749 | 6462859.2642
| 47.6 | 47.5 | 1568360 | 2846350 | 2814640 | 47.5 | 1577956.9521 | 2863499.4837 | 2831157.4235
| 47.8 | 47.7 | 0 | 448043 | 409718 | 47.7 | 0 | 450734.4076 | 412115.1854
| 48 | 47.9 | 0 | 0 | 0 | 47.9 | 0 | 0 | 0
| 48.2 | 48.1 | 0 | 0 | 0 | 48.1 | 0 | 0 | 0
| 48.4 | 48.3 | 0 | 0 | 0 | 48.3 | 0 | 0 | 0
| 48.6 | 48.5 | 0 | 0 | 0 | 48.5 | 0 | 0 | 0
| 48.8 | 48.7 | 0 | 0 | 0 | 48.7 | 0 | 0 | 0
| 49 | 48.9 | 0 | 0 | 0 | 48.9 | 0 | 0 | 0
| 49.2 | 49.1 | 0 | 0 | 0 | 49.1 | 0 | 0 | 0
| 49.4 | 49.3 | 0 | 0 | 0 | 49.3 | 0 | 0 | 0
| 49.6 | 49.5 | 0 | 0 | 0 | 49.5 | 0 | 0 | 0
| 49.8 | 49.7 | 0 | 0 | 0 | 49.7 | 0 | 0 | 0
| 50 | 49.9 | 0 | 0 | 0 | 49.9 | 0 | 0 | 0
| 52.2 | 50.1 | 0 | 0 | 0 | 50.1 | 0 | 0 | 0

## EXP, air, 532 cm
Catalogue of photon fluence spectra emitted by electronic brachytherapy devices
3 June 2021 | Compiled within the EMPIR project "18NRM02 PRISM-eBT" Primary standards and traceable measurement methods for X-ray emitting electronic brachytherapy devices (http://www.ebt-empir.eu/)
This sheet contains spectra of the device: | Papillon50 |  |  | The eBT-equivalent spectrum obtained from this measured spectrum is presented in the sheet "MC, water, 1 cm, eBT-eq."
Manufacturer: | Ariane Medical Systems
Spectra information
origin | measurement + Monte Carlo simulation
detector | Amptek XR-100T CdTe X-ray spectrometer
data evaluation | no deconvolution of fluence spectrum, accuracy of Monte Carlo model tested using the comparison of measured and simulated detector amplitude spectrum
Monte Carlo (MC) code | PENELOPE v.2006
MC simulation settings | N/A, bremstrahlung generation optimization
authors | University of Nice
location | France
publication | Contact radiotherapy using a 50 kV X-ray system: Evaluation of relative dose distribution with the Monte Carlo code PENELOPE and comparison with measurements
DOI | https://doi.org/10.1016/j.radphyschem.2012.01.033
year of the result | 2012
material | air
source-to-detector distance (cm) | 532
description of source-to-detector distance | on device axis, distance from the entrance window of the device
applicator | naked, 22 cm, 25 cm, 30 cm
other information | -
Device setting
nominal bias voltage (kV) | 50
bias voltage sigma (kV) | 48.5
current (mA) | 1.03
characteristics of eBT spectra
Filtration material and density (g/cm3) | Al |  | 2.7 |  | 2.7
Emean (in F; keV) |  |  | 28.7047 |  | 25.7695 | average photon energy
Emean (in Kair; keV) |  |  | 23.3969 |  | 18.1926 | average photon energy weighted by air kerma
Eeff (in Kair; keV) |  |  | 21.3257 |  | 15.5911 | energy of monoenergetic photons with the same HVL as the given spectrum
1st HVL in Al (in Kair; mm) |  |  | 0.89788 |  | 0.36144 | half-value layer - width of the given filtration material required to reduce the Kair to half its original value
Measurement position to applicator end distance (cm): |  |  |  |  | 530.14
recalculation of fluence spectrum from 532 cm to 0 cm
applicator diameter (mm) |  |  | N/A |  | N/A
| Energy bin upper edge (keV) | Energy bin centre (keV) | Fluence (relative) | Energy bin centre (keV) | Fluence (relative)
6.0105
| 6.0961 | 6.0533 | 0.018816 | 6.0533 | 0
| 6.1817 | 6.1389 | 0.018816 | 6.1389 | 0
| 6.2673 | 6.2245 | 0.018816 | 6.2245 | 0
| 6.3529 | 6.3101 | 0.018816 | 6.3101 | 0
| 6.4385 | 6.3957 | 0.018816 | 6.3957 | 0 |  |  | 12891243480.82
| 6.5242 | 6.4814 | 0.018816 | 6.4814 | 0 |  |  | 20820.7802
| 6.6098 | 6.567 | 0.018816 | 6.567 | 0
| 6.6954 | 6.6526 | 0.018816 | 6.6526 | 0 |  | 5.3 | 0
| 6.781 | 6.7382 | 0.018816 | 6.7382 | 0 |  | 5.5 | 0
| 6.8666 | 6.8238 | 0.037632 | 6.8238 | 0 |  | 5.7 | 0
| 6.9522 | 6.9094 | 0.075264 | 6.9094 | 0 |  | 5.9 | 0
| 7.0378 | 6.995 | 0 | 6.995 | 0 |  | 6.1 | 0
| 7.1234 | 7.0806 | 0 | 7.0806 | 0 |  | 6.3 | 0
| 7.209 | 7.1662 | 0 | 7.1662 | 0 |  | 6.5 | 0
| 7.2946 | 7.2518 | 0 | 7.2518 | 0 |  | 6.7 | 0
| 7.3802 | 7.3374 | 0 | 7.3374 | 0 |  | 6.9 | 0
| 7.4659 | 7.423 | 0 | 7.423 | 0 |  | 7.1 | 0
| 7.5515 | 7.5087 | 0 | 7.5087 | 0 |  | 7.3 | 0
| 7.6371 | 7.5943 | 0 | 7.5943 | 0 |  | 7.5 | 0
| 7.7227 | 7.6799 | 0 | 7.6799 | 0 |  | 7.7 | 0
| 7.8083 | 7.7655 | 0 | 7.7655 | 0 |  | 7.9 | 0
| 7.8939 | 7.8511 | 0 | 7.8511 | 0 |  | 8.1 | 0
| 7.9795 | 7.9367 | 0 | 7.9367 | 0 |  | 8.3 | 0
| 8.0651 | 8.0223 | 0 | 8.0223 | 0 |  | 8.5 | 0
| 8.1507 | 8.1079 | 0.23799 | 8.1079 | 105.0344 |  | 8.7 | 0
| 8.2363 | 8.1935 | 0.20042 | 8.1935 | 73.3074 |  | 8.9 | 0
| 8.3219 | 8.2791 | 0.34427 | 8.2791 | 105.2339 |  | 9.1 | 0
| 8.4075 | 8.3647 | 0.36829 | 8.3647 | 94.7472 |  | 9.3 | 0
| 8.4932 | 8.4504 | 1.0335 | 8.4504 | 225.2883 |  | 9.5 | 0
| 8.5788 | 8.536 | 1.0022 | 8.536 | 186.3104 |  | 9.7 | 0.024366
| 8.6644 | 8.6216 | 1.0334 | 8.6216 | 164.8332 |  | 9.9 | 267.1344
| 8.75 | 8.7072 | 1.6613 | 8.7072 | 228.7181 |  | 10.1 | 265.4482
| 8.8356 | 8.7928 | 2.1293 | 8.7928 | 254.5771 |  | 10.3 | 12.2791
| 8.9212 | 8.8784 | 2.6495 | 8.8784 | 276.6911 |  | 10.5 | 14.1269
| 9.0068 | 8.964 | 3.0697 | 8.964 | 281.4137 |  | 10.7 | 15.8536
| 9.0924 | 9.0496 | 3.4975 | 9.0496 | 282.8089 |  | 10.9 | 18.2945
| 9.178 | 9.1352 | 4.6479 | 9.1352 | 333.0165 |  | 11.1 | 22.2523
| 9.2636 | 9.2208 | 5.8417 | 9.2208 | 372.5044 |  | 11.3 | 27.1555
| 9.3492 | 9.3064 | 6.8293 | 9.3064 | 389.1922 |  | 11.5 | 165.9445
| 9.4349 | 9.392 | 11.8544 | 9.392 | 606.4457 |  | 11.7 | 241.333
| 9.5205 | 9.4777 | 10.3165 | 9.4777 | 475.7239 |  | 11.9 | 166.3661
| 9.6061 | 9.5633 | 10.6436 | 9.5633 | 444.0064 |  | 12.1 | 81.4245
| 9.6917 | 9.6489 | 12.8866 | 9.6489 | 488 |  | 12.3 | 97.8689
| 9.7773 | 9.7345 | 15.6679 | 9.7345 | 540.3913 |  | 12.5 | 115.0469
| 9.8629 | 9.8201 | 19.0965 | 9.8201 | 601.8007 |  | 12.7 | 128.2204
| 9.9485 | 9.9057 | 21.9525 | 9.9057 | 634.0295 |  | 12.9 | 148.1116
| 10.0341 | 9.9913 | 26.7613 | 9.9913 | 710.6855 |  | 13.1 | 167.9812
| 10.1197 | 10.0769 | 29.0365 | 10.0769 | 710.9468 |  | 13.3 | 188.8516
| 10.2053 | 10.1625 | 73.6304 | 10.1625 | 1666.5347 |  | 13.5 | 211.812
| 10.2909 | 10.2481 | 53.8455 | 10.2481 | 1129.4849 |  | 13.7 | 233.2542
| 10.3765 | 10.3337 | 59.7805 | 10.3337 | 1165.0155 |  | 13.9 | 255.7996
| 10.4622 | 10.4194 | 65.8063 | 10.4194 | 1194.287 |  | 14.1 | 279.1492
| 10.5478 | 10.505 | 110.1073 | 10.505 | 1865.1538 |  | 14.3 | 302.5295
| 10.6334 | 10.5906 | 86.0735 | 10.5906 | 1364.2562 |  | 14.5 | 324.7971
| 10.719 | 10.6762 | 88.0068 | 10.6762 | 1308.0824 |  | 14.7 | 347.7882
| 10.8046 | 10.7618 | 87.2667 | 10.7618 | 1218.7801 |  | 14.9 | 371.5222
| 10.8902 | 10.8474 | 74.7879 | 10.8474 | 983.3337 |  | 15.1 | 394.0611
| 10.9758 | 10.933 | 28.0534 | 10.933 | 347.8977 |  | 15.3 | 416.7485
| 11.0614 | 11.0186 | 30.3215 | 11.0186 | 355.2936 |  | 15.5 | 441.1884
| 11.147 | 11.1042 | 32.2411 | 11.1042 | 357.5734 |  | 15.7 | 463.9727
| 11.2326 | 11.1898 | 37.0606 | 11.1898 | 389.7667 |  | 15.9 | 487.3256
| 11.3182 | 11.2754 | 42.0605 | 11.2754 | 420.1717 |  | 16.1 | 510.6268
| 11.4039 | 11.361 | 45.2144 | 11.361 | 429.6836 |  | 16.3 | 533.8828
| 11.4895 | 11.4467 | 51.015 | 11.4467 | 461.8755 |  | 16.5 | 557.0548
| 11.5751 | 11.5323 | 56.5869 | 11.5323 | 488.7784 |  | 16.7 | 580.2461
| 11.6607 | 11.6179 | 75.5925 | 11.6179 | 623.7862 |  | 16.9 | 602.3344
| 11.7463 | 11.7035 | 72.1989 | 11.7035 | 569.9291 |  | 17.1 | 626.6951
| 11.8319 | 11.7891 | 75.1205 | 11.7891 | 568.0869 |  | 17.3 | 649.3276
| 11.9175 | 11.8747 | 84.8738 | 11.8747 | 615.669 |  | 17.5 | 670.1335
| 12.0031 | 11.9603 | 93.8528 | 11.9603 | 653.8019 |  | 17.7 | 693.1973
| 12.0887 | 12.0459 | 102.9811 | 12.0459 | 689.7185 |  | 17.9 | 714.0338
| 12.1743 | 12.1315 | 109.3809 | 12.1315 | 705.091 |  | 18.1 | 732.9984
| 12.2599 | 12.2171 | 126.4097 | 12.2171 | 785.1134 |  | 18.3 | 752.8885
| 12.3455 | 12.3027 | 141.8908 | 12.3027 | 849.9591 |  | 18.5 | 771.4381
| 12.4312 | 12.3884 | 109.6662 | 12.3884 | 634.3138 |  | 18.7 | 785.8028
| 12.5168 | 12.474 | 79.3015 | 12.474 | 443.328 |  | 18.9 | 798.236
| 12.6024 | 12.5596 | 91.0592 | 12.5596 | 492.4651 |  | 19.1 | 809.18
| 12.688 | 12.6452 | 91.7726 | 12.6452 | 480.569 |  | 19.3 | 818.0291
| 12.7736 | 12.7308 | 120.5138 | 12.7308 | 611.5628 |  | 19.5 | 833.2596
| 12.8592 | 12.8164 | 90.1624 | 12.8164 | 443.7622 |  | 19.7 | 862.9874
| 12.9448 | 12.902 | 81.4862 | 12.902 | 389.2937 |  | 19.9 | 842.3204
| 13.0304 | 12.9876 | 89.016 | 12.9876 | 413.1643 |  | 20.1 | 835.839
| 13.116 | 13.0732 | 97.8212 | 13.0732 | 441.4481 |  | 20.3 | 840.9411
| 13.2016 | 13.1588 | 110.9505 | 13.1588 | 487.1693 |  | 20.5 | 845.1808
| 13.2872 | 13.2444 | 117.5686 | 13.2444 | 502.6294 |  | 20.7 | 852.2501
| 13.3729 | 13.33 | 126.9366 | 13.33 | 528.7404 |  | 20.9 | 857.8352
| 13.4585 | 13.4157 | 138.3841 | 13.4157 | 561.9855 |  | 21.1 | 864.336
| 13.5441 | 13.5013 | 149.8022 | 13.5013 | 593.4935 |  | 21.3 | 870.8707
| 13.6297 | 13.5869 | 157.6981 | 13.5869 | 609.9546 |  | 21.5 | 876.6673
| 13.7153 | 13.6725 | 167.0862 | 13.6725 | 631.3137 |  | 21.7 | 882.3283
| 13.8009 | 13.7581 | 172.7429 | 13.7581 | 637.9499 |  | 21.9 | 961.5527
| 13.8865 | 13.8437 | 202.6295 | 13.8437 | 731.8326 |  | 22.1 | 1060.5993
| 13.9721 | 13.9293 | 203.5139 | 13.9293 | 719.2164 |  | 22.3 | 991.1302
| 14.0577 | 14.0149 | 210.0618 | 14.0149 | 726.769 |  | 22.5 | 901.5126
| 14.1433 | 14.1005 | 220.9187 | 14.1005 | 748.6621 |  | 22.7 | 905.7102
| 14.2289 | 14.1861 | 218.6796 | 14.1861 | 726.3067 |  | 22.9 | 909.7012
| 14.3145 | 14.2717 | 232.848 | 14.2717 | 758.3148 |  | 23.1 | 913.3853
| 14.4002 | 14.3574 | 250.8173 | 14.3574 | 801.3072 |  | 23.3 | 916.3134
| 14.4858 | 14.443 | 250.6856 | 14.443 | 786.0122 |  | 23.5 | 920.6306
| 14.5714 | 14.5286 | 254.3735 | 14.5286 | 783.1026 |  | 23.7 | 923.4862
| 14.657 | 14.6142 | 277.5924 | 14.6142 | 839.43 |  | 23.9 | 925.4017
| 14.7426 | 14.6998 | 280.2643 | 14.6998 | 832.8241 |  | 24.1 | 928.3266
| 14.8282 | 14.7854 | 284.7802 | 14.7854 | 831.9747 |  | 24.3 | 929.8788
| 14.9138 | 14.871 | 296.7848 | 14.871 | 852.7527 |  | 24.5 | 930.8688
| 14.9994 | 14.9566 | 309.4856 | 14.9566 | 874.9126 |  | 24.7 | 931.2274
| 15.085 | 15.0422 | 316.4475 | 15.0422 | 880.4678 |  | 24.9 | 930.3342
| 15.1706 | 15.1278 | 336.8628 | 15.1278 | 922.7693 |  | 25.1 | 953.2139
| 15.2562 | 15.2134 | 337.6343 | 15.2134 | 910.8882 |  | 25.3 | 956.2745
| 15.3419 | 15.299 | 349.7518 | 15.299 | 929.6192 |  | 25.5 | 920.0007
| 15.4275 | 15.3847 | 357.1465 | 15.3847 | 935.5914 |  | 25.7 | 913.9231
| 15.5131 | 15.4703 | 354.9827 | 15.4703 | 916.8083 |  | 25.9 | 909.5138
| 15.5987 | 15.5559 | 377.242 | 15.5559 | 960.8515 |  | 26.1 | 905.04
| 15.6843 | 15.6415 | 370.0919 | 15.6415 | 929.9094 |  | 26.3 | 899.6794
| 15.7699 | 15.7271 | 385.0318 | 15.7271 | 954.6606 |  | 26.5 | 896.1181
| 15.8555 | 15.8127 | 396.6413 | 15.8127 | 970.7222 |  | 26.7 | 891.6475
| 15.9411 | 15.8983 | 414.6294 | 15.8983 | 1001.9007 |  | 26.9 | 886.3451
| 16.0267 | 15.9839 | 410.5839 | 15.9839 | 979.876 |  | 27.1 | 880.8246
| 16.1123 | 16.0695 | 433.8781 | 16.0695 | 1022.9496 |  | 27.3 | 875.3268
| 16.1979 | 16.1551 | 427.7818 | 16.1551 | 996.6345 |  | 27.5 | 869.8354
| 16.2835 | 16.2407 | 448.3853 | 16.2407 | 1032.5221 |  | 27.7 | 864.2681
| 16.3692 | 16.3264 | 461.5 | 16.3264 | 1050.6513 |  | 27.9 | 858.5458
| 16.4548 | 16.412 | 467.4083 | 16.412 | 1052.2623 |  | 28.1 | 852.5602
| 16.5404 | 16.4976 | 473.467 | 16.4976 | 1054.2918 |  | 28.3 | 846.5617
| 16.626 | 16.5832 | 496.2156 | 16.5832 | 1093.1945 |  | 28.5 | 840.8297
| 16.7116 | 16.6688 | 477.4184 | 16.6688 | 1040.8198 |  | 28.7 | 834.6938
| 16.7972 | 16.7544 | 500.3739 | 16.7544 | 1079.7213 |  | 28.9 | 828.8568
| 16.8828 | 16.84 | 507.825 | 16.84 | 1084.8277 |  | 29.1 | 822.4239
| 16.9684 | 16.9256 | 529.0118 | 16.9256 | 1118.9961 |  | 29.3 | 816.5885
| 17.054 | 17.0112 | 510.3276 | 17.0112 | 1069.0882 |  | 29.5 | 810.464
| 17.1396 | 17.0968 | 533.396 | 17.0968 | 1106.8836 |  | 29.7 | 804.2442
| 17.2252 | 17.1824 | 540.2826 | 17.1824 | 1110.846 |  | 29.9 | 797.8855
| 17.3109 | 17.268 | 532.5304 | 17.268 | 1085.0154 |  | 30.1 | 791.417
| 17.3965 | 17.3537 | 576.3341 | 17.3537 | 1163.8596 |  | 30.3 | 785.0728
| 17.4821 | 17.4393 | 554.0748 | 17.4393 | 1109.1848 |  | 30.5 | 778.8078
| 17.5677 | 17.5249 | 568.1115 | 17.5249 | 1127.5896 |  | 30.7 | 772.4136
| 17.6533 | 17.6105 | 567.6035 | 17.6105 | 1117.1604 |  | 30.9 | 765.916
| 17.7389 | 17.6961 | 593.6448 | 17.6961 | 1158.8394 |  | 31.1 | 759.3215
| 17.8245 | 17.7817 | 615.8665 | 17.7817 | 1192.5771 |  | 31.3 | 752.5655
| 17.9101 | 17.8673 | 595.9215 | 17.8673 | 1144.8742 |  | 31.5 | 745.5866
| 17.9957 | 17.9529 | 628.0028 | 17.9529 | 1197.1902 |  | 31.7 | 738.5011
| 18.0813 | 18.0385 | 625.218 | 18.0385 | 1182.847 |  | 31.9 | 731.54
| 18.1669 | 18.1241 | 609.883 | 18.1241 | 1145.2499 |  | 32.1 | 724.5724
| 18.2525 | 18.2097 | 643.4319 | 18.2097 | 1199.4244 |  | 32.3 | 716.7537
| 18.3382 | 18.2954 | 635.8491 | 18.2954 | 1176.8011 |  | 32.5 | 709.6601
| 18.4238 | 18.381 | 641.0047 | 18.381 | 1178.0224 |  | 32.7 | 703.1851
| 18.5094 | 18.4666 | 667.0084 | 18.4666 | 1217.3692 |  | 32.9 | 696.4824
| 18.595 | 18.5522 | 667.4976 | 18.5522 | 1210.0228 |  | 33.1 | 689.1676
| 18.6806 | 18.6378 | 683.5476 | 18.6378 | 1230.8877 |  | 33.3 | 682.3033
| 18.7662 | 18.7234 | 675.3439 | 18.7234 | 1208.1815 |  | 33.5 | 674.9885
| 18.8518 | 18.809 | 708.366 | 18.809 | 1259.1377 |  | 33.7 | 667.7205
| 18.9374 | 18.8946 | 682.4375 | 18.8946 | 1205.4245 |  | 33.9 | 660.3411
| 19.023 | 18.9802 | 724.0773 | 18.9802 | 1271.0948 |  | 34.1 | 653.5237
| 19.1086 | 19.0658 | 704.3017 | 19.0658 | 1228.8945 |  | 34.3 | 646.0829
| 19.1942 | 19.1514 | 707.4816 | 19.1514 | 1227.0999 |  | 34.5 | 639.1444
| 19.2799 | 19.237 | 741.5009 | 19.237 | 1278.5875 |  | 34.7 | 632.6306
| 19.3655 | 19.3227 | 727.5206 | 19.3227 | 1247.2748 |  | 34.9 | 624.9815
| 19.4511 | 19.4083 | 708.4224 | 19.4083 | 1207.6759 |  | 35.1 | 618.7084
| 19.5367 | 19.4939 | 727.2572 | 19.4939 | 1232.9154 |  | 35.3 | 611.694
| 19.6223 | 19.5795 | 733.1466 | 19.5795 | 1236.144 |  | 35.5 | 604.5083
| 19.7079 | 19.6651 | 760.3546 | 19.6651 | 1275.1684 |  | 35.7 | 597.3776
| 19.7935 | 19.7507 | 726.8997 | 19.7507 | 1212.6581 |  | 35.9 | 589.3408
| 19.8791 | 19.8363 | 755.3119 | 19.8363 | 1253.5486 |  | 36.1 | 582.307
| 19.9647 | 19.9219 | 757.5698 | 19.9219 | 1250.9106 |  | 36.3 | 574.0764
| 20.0503 | 20.0075 | 803.4808 | 20.0075 | 1320.0996 |  | 36.5 | 567.0781
| 20.1359 | 20.0931 | 787.5813 | 20.0931 | 1287.6942 |  | 36.7 | 559.0914
| 20.2215 | 20.1787 | 772.1898 | 20.1787 | 1256.5107 |  | 36.9 | 551.1548
| 20.3072 | 20.2644 | 779.2458 | 20.2644 | 1262.0468 |  | 37.1 | 543.2246
| 20.3928 | 20.35 | 758.3789 | 20.35 | 1222.5863 |  | 37.3 | 534.5191
| 20.4784 | 20.4356 | 789.7828 | 20.4356 | 1267.4359 |  | 37.5 | 526.6406
| 20.564 | 20.5212 | 778.6437 | 20.5212 | 1243.9825 |  | 37.7 | 517.806
| 20.6496 | 20.6068 | 774.5042 | 20.6068 | 1231.9351 |  | 37.9 | 509.4446
| 20.7352 | 20.6924 | 796.8012 | 20.6924 | 1261.934 |  | 38.1 | 500.7407
| 20.8208 | 20.778 | 820.0013 | 20.778 | 1293.173 |  | 38.3 | 491.6428
| 20.9064 | 20.8636 | 826.06 | 20.8636 | 1297.2942 |  | 38.5 | 482.7904
| 20.992 | 20.9492 | 847.7925 | 20.9492 | 1325.9585 |  | 38.7 | 474.0591
| 21.0776 | 21.0348 | 805.4753 | 21.0348 | 1254.6838 |  | 38.9 | 465.9593
| 21.1632 | 21.1204 | 818.2514 | 21.1204 | 1269.5158 |  | 39.1 | 456.5739
| 21.2489 | 21.206 | 818.1761 | 21.206 | 1264.4293 |  | 39.3 | 447.82
| 21.3345 | 21.2917 | 863.1464 | 21.2917 | 1328.7957 |  | 39.5 | 438.6494
| 21.4201 | 21.3773 | 849.6177 | 21.3773 | 1303.0209 |  | 39.7 | 428.9378
| 21.5057 | 21.4629 | 883.3548 | 21.4629 | 1349.716 |  | 39.9 | 419.6686
| 21.5913 | 21.5485 | 833.6241 | 21.5485 | 1269.0587 |  | 40.1 | 409.8245
| 21.6769 | 21.6341 | 842.1477 | 21.6341 | 1277.4042 |  | 40.3 | 400.2986
| 21.7625 | 21.7197 | 856.3726 | 21.7197 | 1294.3608 |  | 40.5 | 390.6241
| 21.8481 | 21.8053 | 889.3571 | 21.8053 | 1339.5061 |  | 40.7 | 380.7057
| 21.9337 | 21.8909 | 822.278 | 21.8909 | 1234.2101 |  | 40.9 | 370.9004
| 22.0193 | 21.9765 | 842.2042 | 21.9765 | 1259.8352 |  | 41.1 | 360.7607
| 22.1049 | 22.0621 | 859.4584 | 22.0621 | 1281.3541 |  | 41.3 | 351.0135
| 22.1905 | 22.1477 | 880.2501 | 22.1477 | 1308.0365 |  | 41.5 | 340.89
| 22.2762 | 22.2334 | 859.5901 | 22.2334 | 1273.1978 |  | 41.7 | 330.6486
| 22.3618 | 22.319 | 834.4143 | 22.319 | 1231.963 |  | 41.9 | 320.3507
| 22.4474 | 22.4046 | 893.6095 | 22.4046 | 1315.2113 |  | 42.1 | 309.7136
| 22.533 | 22.4902 | 871.2937 | 22.4902 | 1278.4005 |  | 42.3 | 299.3235
| 22.6186 | 22.5758 | 826.3611 | 22.5758 | 1208.7809 |  | 42.5 | 288.6703
| 22.7042 | 22.6614 | 889.9404 | 22.6614 | 1297.8758 |  | 42.7 | 277.9104
| 22.7898 | 22.747 | 873.8715 | 22.747 | 1270.6711 |  | 42.9 | 267.2555
| 22.8754 | 22.8326 | 864.7834 | 22.8326 | 1253.7897 |  | 43.1 | 255.9885
| 22.961 | 22.9182 | 872.084 | 22.9182 | 1260.7403 |  | 43.3 | 245.1334
| 23.0466 | 23.0038 | 874.1537 | 23.0038 | 1260.1521 |  | 43.5 | 233.9649
| 23.1322 | 23.0894 | 902.2836 | 23.0894 | 1297.0779 |  | 43.7 | 223.0225
| 23.2179 | 23.175 | 867.0225 | 23.175 | 1242.9658 |  | 43.9 | 211.7765
| 23.3035 | 23.2607 | 853.2115 | 23.2607 | 1219.8548 |  | 44.1 | 200.2252
| 23.3891 | 23.3463 | 882.6209 | 23.3463 | 1258.5333 |  | 44.3 | 189.2764
| 23.4747 | 23.4319 | 877.7852 | 23.4319 | 1248.3433 |  | 44.5 | 177.8931
| 23.5603 | 23.5175 | 849.0532 | 23.5175 | 1204.3476 |  | 44.7 | 166.2869
| 23.6459 | 23.6031 | 905.9716 | 23.6031 | 1281.7943 |  | 44.9 | 153.559
| 23.7315 | 23.6887 | 859.2891 | 23.6887 | 1212.684 |  | 45.1 | 141.3357
| 23.8171 | 23.7743 | 894.0799 | 23.7743 | 1258.6497 |  | 45.3 | 128.4221
| 23.9027 | 23.8599 | 867.6246 | 23.8599 | 1218.4154 |  | 45.5 | 114.4637
| 23.9883 | 23.9455 | 914.6269 | 23.9455 | 1281.3181 |  | 45.7 | 101.2207
| 24.0739 | 24.0311 | 876.167 | 24.0311 | 1224.5135 |  | 45.9 | 86.9382
| 24.1595 | 24.1167 | 870.8233 | 24.1167 | 1214.1837 |  | 46.1 | 73.805
| 24.2452 | 24.2024 | 911.6164 | 24.2024 | 1268.113 |  | 46.3 | 60.6865
| 24.3308 | 24.288 | 919.9331 | 24.288 | 1276.7601 |  | 46.5 | 47.9421
| 24.4164 | 24.3736 | 895.6792 | 24.3736 | 1240.2993 |  | 46.7 | 36.7617
| 24.502 | 24.4592 | 870.6728 | 24.4592 | 1202.9924 |  | 46.9 | 26.4368
| 24.5876 | 24.5448 | 898.5393 | 24.5448 | 1238.7728 |  | 47.1 | 17.4823
| 24.6732 | 24.6304 | 910.6003 | 24.6304 | 1252.6845 |  | 47.3 | 10.3771
| 24.7588 | 24.716 | 879.234 | 24.716 | 1206.9521 |  | 47.5 | 4.546
| 24.8444 | 24.8016 | 883.43 | 24.8016 | 1210.1565 |  | 47.7 | 0.66174
| 24.93 | 24.8872 | 871.162 | 24.8872 | 1190.875 |  | 47.9 | 0
| 25.0156 | 24.9728 | 904.3722 | 24.9728 | 1233.7421 |  | 48.1 | 0
| 25.1012 | 25.0584 | 885.067 | 25.0584 | 1204.9657 |  | 48.3 | 0
| 25.1869 | 25.144 | 890.5989 | 25.144 | 1210.0781 |  | 48.5 | 0
| 25.2725 | 25.2297 | 909.4149 | 25.2297 | 1233.2104 |  | 48.7 | 0
| 25.3581 | 25.3153 | 887.1368 | 25.3153 | 1200.6613 |  | 48.9 | 0
| 25.4437 | 25.4009 | 872.9495 | 25.4009 | 1179.1924 |  | 49.1 | 0
| 25.5293 | 25.4865 | 906.0657 | 25.4865 | 1221.6125 |  | 49.3 | 0
| 25.6149 | 25.5721 | 894.2868 | 25.5721 | 1203.4814 |  | 49.5 | 0
| 25.7005 | 25.6577 | 855.1872 | 25.6577 | 1148.7426 |  | 49.7 | 0
| 25.7861 | 25.7433 | 876.2799 | 25.7433 | 1174.9339 |  | 49.9 | 0
| 25.8717 | 25.8289 | 876.5622 | 25.8289 | 1173.2005 |  | 50.1 | 0
| 25.9573 | 25.9145 | 898.0877 | 25.9145 | 1199.8777
| 26.0429 | 26.0001 | 897.5796 | 26.0001 | 1197.0976
| 26.1285 | 26.0857 | 863.0711 | 26.0857 | 1149.0867
| 26.2142 | 26.1714 | 892.4052 | 26.1714 | 1186.1163
| 26.2998 | 26.257 | 858.1601 | 26.257 | 1138.6796
| 26.3854 | 26.3426 | 888.2469 | 26.3426 | 1176.6409
| 26.471 | 26.4282 | 892.7816 | 26.4282 | 1180.7047
| 26.5566 | 26.5138 | 908.3988 | 26.5138 | 1199.4088
| 26.6422 | 26.5994 | 893.0826 | 26.5994 | 1177.2957
| 26.7278 | 26.685 | 858.6305 | 26.685 | 1130.0913
| 26.8134 | 26.7706 | 863.9366 | 26.7706 | 1135.3
| 26.899 | 26.8562 | 863.4098 | 26.8562 | 1132.8579
| 26.9846 | 26.9418 | 910.3369 | 26.9418 | 1192.6098
| 27.0702 | 27.0274 | 868.9229 | 27.0274 | 1136.6405
| 27.1559 | 27.113 | 876.9197 | 27.113 | 1145.3947
| 27.2415 | 27.1987 | 848.7709 | 27.1987 | 1106.9988
| 27.3271 | 27.2843 | 856.8054 | 27.2843 | 1115.8583
| 27.4127 | 27.3699 | 855.206 | 27.3699 | 1112.1805
| 27.4983 | 27.4555 | 827.8476 | 27.4555 | 1075.0777
| 27.5839 | 27.5411 | 841.9596 | 27.5411 | 1091.8749
| 27.6695 | 27.6267 | 861.3212 | 27.6267 | 1115.4396
| 27.7551 | 27.7123 | 852.8916 | 27.7123 | 1103.0141
| 27.8407 | 27.7979 | 867.7751 | 27.7979 | 1120.7474
| 27.9263 | 27.8835 | 861.0013 | 27.8835 | 1110.5184
| 28.0119 | 27.9691 | 875.452 | 27.9691 | 1127.6706
| 28.0975 | 28.0547 | 838.2904 | 28.0547 | 1078.3978
| 28.1832 | 28.1404 | 844.0293 | 28.1404 | 1084.384
| 28.2688 | 28.226 | 837.2556 | 28.226 | 1074.3136
| 28.3544 | 28.3116 | 856.6172 | 28.3116 | 1097.7757
| 28.44 | 28.3972 | 854.1523 | 28.3972 | 1093.2573
| 28.5256 | 28.4828 | 828.5626 | 28.4828 | 1059.2042
| 28.6112 | 28.5684 | 792.2665 | 28.5684 | 1011.5772
| 28.6968 | 28.654 | 865.4984 | 28.654 | 1103.7566
| 28.7824 | 28.7396 | 805.0614 | 28.7396 | 1025.4658
| 28.868 | 28.8252 | 833.1725 | 28.8252 | 1060.0297
| 28.9536 | 28.9108 | 845.7039 | 28.9108 | 1074.7268
| 29.0392 | 28.9964 | 797.7796 | 28.9964 | 1012.6635
| 29.1249 | 29.082 | 828.8072 | 29.082 | 1050.8594
| 29.2105 | 29.1677 | 792.5299 | 29.1677 | 1003.7397
| 29.2961 | 29.2533 | 790.1779 | 29.2533 | 999.6547
| 29.3817 | 29.3389 | 800.3197 | 29.3389 | 1011.3783
| 29.4673 | 29.4245 | 848.0183 | 29.4245 | 1070.4971
| 29.5529 | 29.5101 | 844.7255 | 29.5101 | 1065.2002
| 29.6385 | 29.5957 | 810.8379 | 29.5957 | 1021.3872
| 29.7241 | 29.6813 | 819.5685 | 29.6813 | 1031.3074
| 29.8097 | 29.7669 | 815.8429 | 29.7669 | 1025.5594
| 29.8953 | 29.8525 | 802.4648 | 29.8525 | 1007.7121
| 29.9809 | 29.9381 | 818.1949 | 29.9381 | 1026.4277
| 30.0665 | 30.0237 | 791.7585 | 30.0237 | 992.2784
| 30.1522 | 30.1094 | 793.132 | 30.1094 | 993.0455
| 30.2378 | 30.195 | 793.7718 | 30.195 | 992.9032
| 30.3234 | 30.2806 | 816.1252 | 30.2806 | 1019.9068
| 30.409 | 30.3662 | 789.8016 | 30.3662 | 986.0944
| 30.4946 | 30.4518 | 777.5148 | 30.4518 | 969.8622
| 30.5802 | 30.5374 | 783.7052 | 30.5374 | 976.6956
| 30.6658 | 30.623 | 806.3032 | 30.623 | 1003.9547
| 30.7514 | 30.7086 | 750.7584 | 30.7086 | 933.9619
| 30.837 | 30.7942 | 814.8833 | 30.7942 | 1012.8425
| 30.9226 | 30.8798 | 800.7525 | 30.8798 | 994.4129
| 31.0082 | 30.9654 | 758.7364 | 30.9654 | 941.4237
| 31.0939 | 31.051 | 786.9039 | 31.051 | 975.5411
| 31.1795 | 31.1367 | 805.513 | 31.1367 | 997.7684
| 31.2651 | 31.2223 | 785.1729 | 31.2223 | 971.7612
| 31.3507 | 31.3079 | 784.8906 | 31.3079 | 970.6087
| 31.4363 | 31.3935 | 762.5937 | 31.3935 | 942.2647
| 31.5219 | 31.4791 | 803.1798 | 31.4791 | 991.6105
| 31.6075 | 31.5647 | 730.2113 | 31.5647 | 900.8014
| 31.6931 | 31.6503 | 745.076 | 31.6503 | 918.4101
| 31.7787 | 31.7359 | 753.3926 | 31.7359 | 927.9327
| 31.8643 | 31.8215 | 760.9943 | 31.8215 | 936.5671
| 31.9499 | 31.9071 | 748.35 | 31.9071 | 920.2969
| 32.0355 | 31.9927 | 712.0539 | 31.9927 | 874.9947
| 32.1212 | 32.0784 | 717.2659 | 32.0784 | 880.7358
| 32.2068 | 32.164 | 723.193 | 32.164 | 887.3516
| 32.2924 | 32.2496 | 738.3022 | 32.2496 | 905.2216
| 32.378 | 32.3352 | 711.0566 | 32.3352 | 871.1788
| 32.4636 | 32.4208 | 734.5014 | 32.4208 | 899.2514
| 32.5492 | 32.5064 | 742.9121 | 32.5064 | 908.8964
| 32.6348 | 32.592 | 741.7644 | 32.592 | 906.8484
| 32.7204 | 32.6776 | 737.5872 | 32.6776 | 901.1085
| 32.806 | 32.7632 | 726.3164 | 32.7632 | 886.7219
| 32.8916 | 32.8488 | 732.7515 | 32.8488 | 893.962
| 32.9772 | 32.9344 | 704.9414 | 32.9344 | 859.4468
| 33.0629 | 33.02 | 710.0782 | 33.02 | 865.1243
| 33.1485 | 33.1057 | 718.6207 | 33.1057 | 874.9459
| 33.2341 | 33.1913 | 730.5312 | 33.1913 | 888.8583
| 33.3197 | 33.2769 | 694.3104 | 33.2769 | 844.2335
| 33.4053 | 33.3625 | 717.0778 | 33.3625 | 871.3506
| 33.4909 | 33.4481 | 714.1048 | 33.4481 | 867.1796
| 33.5765 | 33.5337 | 699.6541 | 33.5337 | 849.0896
| 33.6621 | 33.6193 | 708.0649 | 33.6193 | 858.754
| 33.7477 | 33.7049 | 724.6606 | 33.7049 | 878.3315
| 33.8333 | 33.7905 | 689.8322 | 33.7905 | 835.5996
| 33.9189 | 33.8761 | 707.5192 | 33.8761 | 856.4985
| 34.0045 | 33.9617 | 691.5068 | 33.9617 | 836.6057
| 34.0902 | 34.0474 | 673.5375 | 34.0474 | 814.3753
| 34.1758 | 34.133 | 695.9474 | 34.133 | 840.969
| 34.2614 | 34.2186 | 667.2906 | 34.2186 | 805.864
| 34.347 | 34.3042 | 659.3503 | 34.3042 | 795.8081
| 34.4326 | 34.3898 | 652.8588 | 34.3898 | 787.5162
| 34.5182 | 34.4754 | 649.7541 | 34.4754 | 783.321
| 34.6038 | 34.561 | 676.3788 | 34.561 | 814.9546
| 34.6894 | 34.6466 | 656.7725 | 34.6466 | 790.885
| 34.775 | 34.7322 | 700.896 | 34.7322 | 843.5466
| 34.8606 | 34.8178 | 654.7027 | 34.8178 | 787.5151
| 34.9462 | 34.9034 | 650.9395 | 34.9034 | 782.5584
| 35.0319 | 34.989 | 669.9813 | 34.989 | 805.0126
| 35.1175 | 35.0747 | 637.3544 | 35.0747 | 765.3973
| 35.2031 | 35.1603 | 656.3397 | 35.1603 | 787.7759
| 35.2887 | 35.2459 | 672.3709 | 35.2459 | 806.5904
| 35.3743 | 35.3315 | 564.7622 | 35.3315 | 677.1453
| 35.4599 | 35.4171 | 640.2332 | 35.4171 | 767.2351
| 35.5455 | 35.5027 | 638.9349 | 35.5027 | 765.2845
| 35.6311 | 35.5883 | 647.9478 | 35.5883 | 775.6836
| 35.7167 | 35.6739 | 636.4512 | 35.6739 | 761.5353
| 35.8023 | 35.7595 | 641.1552 | 35.7595 | 766.7791
| 35.8879 | 35.8451 | 612.3103 | 35.8451 | 731.9184
| 35.9735 | 35.9307 | 583.1455 | 35.9307 | 696.713
| 36.0592 | 36.0164 | 611.313 | 36.0164 | 730.0093
| 36.1448 | 36.102 | 642.9427 | 36.102 | 767.4084
| 36.2304 | 36.1876 | 610.598 | 36.1876 | 728.4526
| 36.316 | 36.2732 | 595.263 | 36.2732 | 709.82
| 36.4016 | 36.3588 | 614.1542 | 36.3588 | 732.0016
| 36.4872 | 36.4444 | 607.4746 | 36.4444 | 723.7017
| 36.5728 | 36.53 | 614.5117 | 36.53 | 731.7459
| 36.6584 | 36.6156 | 621.5489 | 36.6156 | 739.7854
| 36.744 | 36.7012 | 586.7205 | 36.7012 | 698.0133
| 36.8296 | 36.7868 | 556.6714 | 36.7868 | 661.9654
| 36.9152 | 36.8724 | 587.8871 | 36.8724 | 698.7727
| 37.0009 | 36.958 | 571.6489 | 36.958 | 679.1701
| 37.0865 | 37.0437 | 556.1069 | 37.0437 | 660.414
| 37.1721 | 37.1293 | 580.9816 | 37.1293 | 689.6533
| 37.2577 | 37.2149 | 554.1688 | 37.2149 | 657.5405
| 37.3433 | 37.3005 | 596.185 | 37.3005 | 707.0903
| 37.4289 | 37.3861 | 533.5653 | 37.3861 | 632.5528
| 37.5145 | 37.4717 | 527.2431 | 37.4717 | 624.7942
| 37.6001 | 37.5573 | 533.5653 | 37.5573 | 632.0217
| 37.6857 | 37.6429 | 547.9407 | 37.6429 | 648.7806
| 37.7713 | 37.7285 | 528.3345 | 37.7285 | 625.3087
| 37.8569 | 37.8141 | 540.452 | 37.8141 | 639.3893
| 37.9425 | 37.8997 | 532.8691 | 37.8997 | 630.1631
| 38.0282 | 37.9854 | 519.8296 | 37.9854 | 614.4964
| 38.1138 | 38.071 | 539.8875 | 38.071 | 637.9532
| 38.1994 | 38.1566 | 537.9494 | 38.1566 | 635.4124
| 38.285 | 38.2422 | 541.6374 | 38.2422 | 639.5182
| 38.3706 | 38.3278 | 483.4395 | 38.3278 | 570.5816
| 38.4562 | 38.4134 | 530.0844 | 38.4134 | 625.3935
| 38.5418 | 38.499 | 514.2601 | 38.499 | 606.4923
| 38.6274 | 38.5846 | 530.63 | 38.5846 | 625.5613
| 38.713 | 38.6702 | 545.2877 | 38.6702 | 642.6
| 38.7986 | 38.7558 | 517.9668 | 38.7558 | 610.1761
| 38.8842 | 38.8414 | 475.4051 | 38.8414 | 559.8305
| 38.9699 | 38.927 | 494.7667 | 38.927 | 582.4169
| 39.0555 | 39.0127 | 499.8847 | 39.0127 | 588.2275
| 39.1411 | 39.0983 | 495.3688 | 39.0983 | 582.7032
| 39.2267 | 39.1839 | 489.724 | 39.1839 | 575.8573
| 39.3123 | 39.2695 | 491.0976 | 39.2695 | 577.2676
| 39.3979 | 39.3551 | 447.2187 | 39.3551 | 525.5045
| 39.4835 | 39.4407 | 481.3321 | 39.4407 | 565.392
| 39.5691 | 39.5263 | 489.5735 | 39.5263 | 574.8734
| 39.6547 | 39.6119 | 478.679 | 39.6119 | 561.8874
| 39.7403 | 39.6975 | 488.087 | 39.6975 | 572.7354
| 39.8259 | 39.7831 | 491.8502 | 39.7831 | 576.9561
| 39.9115 | 39.8687 | 472.6015 | 39.8687 | 554.1907
| 39.9972 | 39.9544 | 456.4573 | 39.9544 | 535.0813
| 40.0828 | 40.04 | 482.2917 | 40.04 | 565.1799
| 40.1684 | 40.1256 | 458.8282 | 40.1256 | 537.5099
| 40.254 | 40.2112 | 452.2237 | 40.2112 | 529.6028
| 40.3396 | 40.2968 | 448.9121 | 40.2968 | 525.557
| 40.4252 | 40.3824 | 434.1604 | 40.3824 | 508.1261
| 40.5108 | 40.468 | 456.0434 | 40.468 | 533.5699
| 40.5964 | 40.5536 | 409.0787 | 40.5536 | 478.4724
| 40.682 | 40.6392 | 424.3008 | 40.6392 | 496.1235
| 40.7676 | 40.7248 | 440.7084 | 40.7248 | 515.1504
| 40.8532 | 40.8104 | 410.8285 | 40.8104 | 480.0774
| 40.9389 | 40.896 | 461.5941 | 40.896 | 539.2371
| 41.0245 | 40.9817 | 433.5395 | 40.9817 | 506.3119
| 41.1101 | 41.0673 | 413.8203 | 41.0673 | 483.1392
| 41.1957 | 41.1529 | 411.8446 | 41.1529 | 480.6907
| 41.2813 | 41.2385 | 417.3201 | 41.2385 | 486.9389
| 41.3669 | 41.3241 | 380.4783 | 41.3241 | 443.8221
| 41.4525 | 41.4097 | 401.2136 | 41.4097 | 467.8743
| 41.5381 | 41.4953 | 406.6702 | 41.4953 | 474.1018
| 41.6237 | 41.5809 | 387.2521 | 41.5809 | 451.3357
| 41.7093 | 41.6665 | 378.258 | 41.6665 | 440.729
| 41.7949 | 41.7521 | 413.8391 | 41.7521 | 482.0515
| 41.8805 | 41.8377 | 393.1227 | 41.8377 | 457.7932
| 41.9662 | 41.9234 | 374.5513 | 41.9234 | 436.0464
| 42.0518 | 42.009 | 372.8579 | 42.009 | 433.956
| 42.1374 | 42.0946 | 435.835 | 42.0946 | 507.1149
| 42.223 | 42.1802 | 373.0648 | 42.1802 | 433.9618
| 42.3086 | 42.2658 | 398.7675 | 42.2658 | 463.7358
| 42.3942 | 42.3514 | 383.3572 | 42.3514 | 445.6962
| 42.4798 | 42.437 | 366.6674 | 42.437 | 426.1799
| 42.5654 | 42.5226 | 359.4044 | 42.5226 | 417.6286
| 42.651 | 42.6082 | 358.3131 | 42.6082 | 416.2521
| 42.7366 | 42.6938 | 369.8661 | 42.6938 | 429.5622
| 42.8222 | 42.7794 | 378.879 | 42.7794 | 439.9169
| 42.9079 | 42.865 | 348.2089 | 42.865 | 404.2029
| 42.9935 | 42.9507 | 337.1639 | 42.9507 | 391.2829
| 43.0791 | 43.0363 | 338.3305 | 43.0363 | 392.5381
| 43.1647 | 43.1219 | 341.7362 | 43.1219 | 396.3906
| 43.2503 | 43.2075 | 332.0648 | 43.2075 | 385.077
| 43.3359 | 43.2931 | 320.5117 | 43.2931 | 371.5883
| 43.4215 | 43.3787 | 301.9215 | 43.3787 | 349.9501
| 43.5071 | 43.4643 | 313.0042 | 43.4643 | 362.7079
| 43.5927 | 43.5499 | 314.66 | 43.5499 | 364.5388
| 43.6783 | 43.6355 | 299.4755 | 43.6355 | 346.8644
| 43.7639 | 43.7211 | 312.6843 | 43.7211 | 362.0774
| 43.8495 | 43.8067 | 288.5245 | 43.8067 | 334.0225
| 43.9352 | 43.8924 | 300.7361 | 43.8924 | 348.0783
| 44.0208 | 43.978 | 268.9371 | 43.978 | 311.2011
| 44.1064 | 44.0636 | 306.8325 | 44.0636 | 354.9699
| 44.192 | 44.1492 | 314.7352 | 44.1492 | 364.0289
| 44.2776 | 44.2348 | 301.1689 | 44.2348 | 348.2585
| 44.3632 | 44.3204 | 272.418 | 44.3204 | 314.941
| 44.4488 | 44.406 | 303.4456 | 44.406 | 350.733
| 44.5344 | 44.4916 | 293.5108 | 44.4916 | 339.1743
| 44.62 | 44.5772 | 268.636 | 44.5772 | 310.3608
| 44.7056 | 44.6628 | 282.447 | 44.6628 | 326.2451
| 44.7912 | 44.7484 | 271.9288 | 44.7484 | 314.0273
| 44.8769 | 44.834 | 244.5139 | 44.834 | 282.3068
| 44.9625 | 44.9197 | 233.9393 | 44.9197 | 270.0395
| 45.0481 | 45.0053 | 259.1904 | 45.0053 | 299.123
| 45.1337 | 45.0909 | 242.8769 | 45.0909 | 280.2365
| 45.2193 | 45.1765 | 239.3019 | 45.1765 | 276.0532
| 45.3049 | 45.2621 | 256.2175 | 45.2621 | 295.5046
| 45.3905 | 45.3477 | 252.078 | 45.3477 | 290.6696
| 45.4761 | 45.4333 | 235.2188 | 45.4333 | 271.1732
| 45.5617 | 45.5189 | 230.703 | 45.5189 | 265.9123
| 45.6473 | 45.6045 | 232.0577 | 45.6045 | 267.4191
| 45.7329 | 45.6901 | 222.7814 | 45.6901 | 256.6771
| 45.8185 | 45.7757 | 219.5263 | 45.7757 | 252.8756
| 45.9042 | 45.8614 | 209.1022 | 45.8614 | 240.8196
| 45.9898 | 45.947 | 200.6726 | 45.947 | 231.0654
| 46.0754 | 46.0326 | 201.1242 | 46.0326 | 231.5395
| 46.161 | 46.1182 | 187.3697 | 46.1182 | 215.6625
| 46.2466 | 46.2038 | 183.9264 | 46.2038 | 211.6578
| 46.3322 | 46.2894 | 194.5386 | 46.2894 | 223.8266
| 46.4178 | 46.375 | 186.6735 | 46.375 | 214.736
| 46.5034 | 46.4606 | 176.7011 | 46.4606 | 203.2254
| 46.589 | 46.5462 | 159.1457 | 46.5462 | 183
| 46.6746 | 46.6318 | 162.9277 | 46.6318 | 187.3134
| 46.7602 | 46.7174 | 145.2407 | 46.7174 | 166.9477
| 46.8459 | 46.803 | 148.8722 | 46.803 | 171.0899
| 46.9315 | 46.8887 | 159.5973 | 46.8887 | 183.3815
| 47.0171 | 46.9743 | 160.143 | 46.9743 | 183.9746
| 47.1027 | 47.0599 | 153.5762 | 47.0599 | 176.3982
| 47.1883 | 47.1455 | 147.2164 | 47.1455 | 169.0624
| 47.2739 | 47.2311 | 145.5794 | 47.2311 | 167.1522
| 47.3595 | 47.3167 | 127.6289 | 47.3167 | 146.5153
| 47.4451 | 47.4023 | 148.3453 | 47.4023 | 170.2669
| 47.5307 | 47.4879 | 137.4509 | 47.4879 | 157.7344
| 47.6163 | 47.5735 | 117.6376 | 47.5735 | 134.9735
| 47.7019 | 47.6591 | 117.8634 | 47.6591 | 135.2088
| 47.7875 | 47.7447 | 108.1167 | 47.7447 | 124.0061
| 47.8732 | 47.8304 | 82.4517 | 47.8304 | 94.5528
| 47.9588 | 47.916 | 87.5697 | 47.916 | 100.4045
| 48.0444 | 48.0016 | 92.1043 | 48.0016 | 105.5857
| 48.13 | 48.0872 | 77.5972 | 48.0872 | 88.94
| 48.2156 | 48.1728 | 72.9308 | 48.1728 | 83.5774
| 48.3012 | 48.2584 | 62.1492 | 48.2584 | 71.2099
| 48.3868 | 48.344 | 55.94 | 48.344 | 64.0846
| 48.4724 | 48.4296 | 51.8381 | 48.4296 | 59.3757
| 48.558 | 48.5152 | 42.543 | 48.5152 | 48.7209
| 48.6404 | 48.6008 | 34.3956 | 48.6008 | 39.384
| 48.72 | 48.68 | 30.7077 | 48.68 | 35.1559
| 48.8 | 48.76 | 0 | 48.76 | 0